## Sheet1
| Web Address | Report Number | Event Date | Date Received | Event Type | Device Problem | Brand Name | Patient Problem | PMA/PMN Number | Number of Events | Event Text
| outpatient diagnostic or treatment facilities | nursing homes and ambulatory surgical facilities. Device user facilities must submit reports when they become aware of information that reasonably suggests that a device may have caused or contributed to a death or serious injury of a patient in their facility. Death reports must be sent to the FDA and the manufacturer | if known. Serious injury reports must be sent to the manufacturer or to the FDA |  | if the manufacturer is not known.
| product | and company names affect search results. Searches only retrieve records that contain the search term(s) provided by the requester.Submission of a medical device report and the FDA's release of that information is not necessarily an admission that a product | user facility | manufacturer | importer
| importers and device user facilities) and voluntary reporters such as health care professionals | patients and consumers. Each year | the FDA receives several hundred thousand medical device reports (MDRs) of suspected device-associated deaths | and contribute to benefit-risk assessments of these products. The MAUDE database houses MDRs submitted to the FDA by mandatory reporters (manufacturers | serious injuries and malfunctions. The FDA uses MDRs to monitor device performance |  | patients and consumers.
| this passive surveillance system has limitations | including the potential submission of incomplete | inaccurate | or biased data. In addition | untimely |  | MDRs comprise only one of the FDA's several important postmarket surveillance data sources.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21269899&pc=LWP | 2124215-2025-04594 | 2024-12-12 05:00 | 2025-01-29 05:00 | Death | Pacing Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED BY THE FACILITY THAT FOLLOWING IMPLANT OF THE PACEMAKER THE PATIENT CODED AND EXPIRED. THERE WAS NOTED TO BE AN ABSENCE OF PACEMAKER SPIKES DURING THE CODE. THE PACEMAKER WAS PROGRAMMED TO DDDR MODE WITH A LOWER RATE OF 60 BEATS/MINUTE AND AN UPPER RATE OF 120 BEATS/MINUTE. IT WAS ALSO REPORTED THAT THE PATIENT HAD BEEN CRITICALLY ILL. NO OTHER INFORMATION (INCLUDING IDENTIFIABLE INFORMATION) WAS PROVIDED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21218816&pc=LWP | MW5165117 |  | 2025-01-22 05:00 | Injury |  | ACCOLADE MRI DDDR | Tachycardia |  | 1 | Event Description: IT WAS REPORTED THROUGH A RESEARCH ARTICLE THAT THE PATIENT PRESENTED IN CLINIC WITH EPISODE(S) OF PACEMAKER MEDIATED TACHYCARDIA (PMT). IT WAS NOTED THAT THIS WAS THE FIRST EPISODE OF PMT AFTER 4 YEARS OF PACEMAKER IMPLANT. PROGRAMMING CHANGES WERE COMPLETED ON THE PULSE GENERATOR DEVICE. SPECIFIC DEVICE AND/OR PATIENT INFORMATION ARE NOT MADE AVAILABLE IN THE ARTICLE. DETAILS ARE LISTED IN THE ATTACHED ARTICLE, TITLED ¿ENDLESS LOOP TACHYCARDIA AMONG PATIENTS WITH DEVICES HAVING ADVANCED PREVENTIVE ALGORITHMS: A CASE SERIES AND BRIEF REVIEW." THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2).
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21289208&pc=LWP | 2124215-2025-05210 | 2025-01-07 05:00 | 2025-01-31 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO INFECTION. THERE WAS NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21289297&pc=LWP | 2124215-2025-05218 | 2025-01-03 05:00 | 2025-01-31 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF THE SYSTEM REVISION DUE TO INFECTION. NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THE PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21277017&pc=LWP | 2124215-2025-04849 | 2025-01-07 05:00 | 2025-01-30 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Endocarditis; Unspecified Infection | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PRODUCT WAS PART OF A SYSTEM REVISION DUE TO INFECTION. IT WAS NOTED THAT THE INFECTION RESULTED IN ENDOCARDITIS. THERE WERE NO ADDITIONAL ADVERSE EFFECTS REPORTED. Manufacturer Narrative: AS NO FURTHER INFORMATION CONCERNING THIS REPORT IS EXPECTED, OUR INVESTIGATION IS COMPLETE. THIS INVESTIGATION WILL BE UPDATED SHOULD FURTHER INFORMATION BE PROVIDED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21261962&pc=LWP | 2124215-2025-04344 | 2024-12-31 05:00 | 2025-01-29 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Unspecified Infection; Pocket Erosion | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF THE SYSTEM REVISION DUE TO POCKET EROSION AND INFECTION. REPORTEDLY, THE PATIENT'S IMPLANTED DEVICE WAS ERODED OUT OF THE BODY. THE PATIENT WAS TREATED WITH INTRAVENOUS ANTIBIOTICS. NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THE PACEMAKER REMAINS IMPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21262423&pc=LWP | 2124215-2025-04354 | 2024-12-31 05:00 | 2025-01-29 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Endocarditis; Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF THE SYSTEM REVISION DUE TO INFECTION AND ENDOCARDITIS. REPORTEDLY, A MITRAL VALVE REPLACEMENT (MVR) WAS PERFORMED. NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THE PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21262752&pc=LWP | 2124215-2025-04367 | 2025-01-06 05:00 | 2025-01-29 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection; Pocket Erosion | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO INFECTION AND POCKET EROSION. THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THIS PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21262762&pc=LWP | 2124215-2025-04361 | 2025-01-06 05:00 | 2025-01-29 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO INFECTION. THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THIS PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21264131&pc=LWP | 2124215-2025-04357 | 2025-01-06 05:00 | 2025-01-29 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO INFECTION. ACCORDING TO REPORTS, THE PATIENT HAS SOME BRAIN ISSUES. THIS DEVICE WILL SOON BE REMOVED. THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21264133&pc=LWP | 2124215-2025-04392 | 2025-01-07 05:00 | 2025-01-29 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Bacterial Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO A BACTEREMIA INFECTION. THERE WAS NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21264168&pc=LWP | 2124215-2025-04333 | 2024-12-06 05:00 | 2025-01-29 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Hematoma; Unspecified Infection; Pocket Erosion | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS IMPLANTABLE PACEMAKER DEVICE EXPERIENCED HEMATOMA AND HAD AN EROSION IN THE POCKET OF THE DEVICE. THE PATIENT THEN HAD A POCKET REVISION. THE IMPLANTABLE PACEMAKER DEVICE REMAINS IN SERVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21264650&pc=LWP | 2124215-2025-04414 | 2025-01-14 05:00 | 2025-01-29 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS PACEMAKER SYSTEM UNDERWENT A REVISION SURGERY IN WHICH THIS COMPLETE SYSTEM WAS EXPLANTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER HAS NOT BEEN RECEIVED FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21270176&pc=LWP | 2124215-2025-04623 | 2025-01-07 05:00 | 2025-01-29 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO INFECTION. THERE WAS NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21270177&pc=LWP | 2124215-2025-04632 | 2025-01-07 05:00 | 2025-01-29 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE EL DR | Unspecified Infection | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO INFECTION. THERE WAS NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21250182&pc=LWP | 2124215-2025-04032 | 2024-12-30 05:00 | 2025-01-28 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF THE SYSTEM REVISION DUE TO INFECTION. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THE PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21251283&pc=LWP | 2124215-2025-04039 | 2025-01-02 05:00 | 2025-01-28 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO INFECTION. THERE WAS NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21251653&pc=LWP | 2124215-2025-04044 | 2025-01-03 05:00 | 2025-01-28 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO INFECTION. THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THE PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21255255&pc=LWP | 2124215-2025-04152 | 2025-01-09 05:00 | 2025-01-28 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE EL DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXPLANTED DUE TO AN UNKNOWN PRODUCT PERFORMANCE ISSUE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. AT THIS TIME, THIS DEVICE WAS ALREADY RETURNED TO BOSTON SCIENTIFIC, BUT FOR FURTHER ANALYSIS HAS NOT YET BEEN CONCLUDED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21256493&pc=LWP | 2124215-2025-04167 | 2025-01-17 05:00 | 2025-01-28 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection |  | 1 | Event Description: IT WAS REPORTED THAT THIS PRODUCT WAS PART OF A SYSTEM REVISION DUE TO INFECTION. THERE WERE NO ADDITIONAL ADVERSE EFFECTS REPORTED. Manufacturer Narrative: AS NO FURTHER INFORMATION CONCERNING THIS REPORT IS EXPECTED, OUR INVESTIGATION IS COMPLETE. THIS INVESTIGATION WILL BE UPDATED SHOULD FURTHER INFORMATION BE PROVIDED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21258665&pc=LWP | 2124215-2025-04253 | 2025-01-06 05:00 | 2025-01-28 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO INFECTION. THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THIS PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21242845&pc=LWP | 2124215-2024-78215 | 2024-10-18 04:00 | 2025-01-27 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Unspecified Infection; Pocket Erosion | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF THE SYSTEM REVISION DUE TO INFECTION AND EROSION. THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THE PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21248874&pc=LWP | 2124215-2025-03943 | 2025-01-13 05:00 | 2025-01-27 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Syncope/Fainting | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS PACEMAKER HAD A SYNCOPAL EPISODE. TECHNICAL SERVICES (TS) WAS CONSULTED AND PATIENT INITIATED INTERROGATION WAS UNSUCCESSFUL. TS REFERRED THE PATIENT TO THEIR CLINIC FOR FURTHER EXAMINATION. THIS DEVICE REMAINS IN SERVICE. NO FURTHER INFORMATION IS AVAILABLE FROM THE FIELD AT THIS TIME. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21249102&pc=LWP | 2124215-2025-04007 | 2024-12-30 05:00 | 2025-01-27 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Bacterial Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF THE SYSTEM REVISION DUE TO BACTERIAL INFECTION. REPORTEDLY, A MASS GREW ON ONE OF THE IMPLANTED PRODUCTS. NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THE PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21249258&pc=LWP | 2124215-2025-04021 | 2024-12-29 05:00 | 2025-01-27 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Unspecified Infection; Pain; Fluid Discharge | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF THE SYSTEM REVISION DUE TO INFECTION. REPORTEDLY, THE PATIENT HAD POCKET PAIN AND FLUID DISCHARGE CAME OUT AFTER THE POCKET WAS OPENED. NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THE PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21237235&pc=LWP | 2124215-2025-03634 | 2025-01-03 05:00 | 2025-01-26 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO INFECTION. THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THE PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21231273&pc=LWP | 2124215-2025-03437 | 2024-12-30 05:00 | 2025-01-24 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Erosion; Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO INFECTION AND EROSION. THERE WAS NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21233634&pc=LWP | 2124215-2025-03552 | 2025-01-14 05:00 | 2025-01-24 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection |  | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS PACEMAKER SYSTEM EXPERIENCED AN INFECTION. SUBSEQUENTLY, THE PATIENT UNDERWENT A SURGICAL PROCEDURE, AND THIS SYSTEM WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED OUTSIDE OF THE SURGICAL PROCEDURE. THE EXPLANTED DEVICE AND BOTH NON-BOSTON SCIENTIFIC LEADS WERE DISPOSED OF AND WILL NOT RETURN FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21209683&pc=LWP | 2124215-2025-02922 | 2024-06-06 04:00 | 2025-01-22 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PRODUCT WAS PART OF A SYSTEM REVISION DUE TO INFECTION. THIS DEVICE WAS EXPLANTED. A NEW SYSTEM WAS SUCCESSFULLY IMPLANTED. THERE WERE NO ADDITIONAL ADVERSE EFFECTS REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21209688&pc=LWP | 2124215-2025-02916 | 2024-07-11 04:00 | 2025-01-22 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PRODUCT WAS PART OF A SYSTEM REVISION DUE TO INFECTION. THIS DEVICE WAS EXPLANTED. AVAILABLE INFORMATION DOES NOT INDICATE THAT A NEW SYSTEM WAS IMPLANTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21210385&pc=LWP | 2124215-2025-02963 | 2024-09-05 04:00 | 2025-01-22 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Pocket Erosion | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS UNKNOWN DEVICE WAS EXPLANTED DUE TO EROSION WAS IMMINENT. THIS IS ALL THE INFORMATION RECEIVED, AND ADDITIONAL INFORMATION HAS BEEN REQUESTED. ADDITIONAL INFORMATION WAS RECEIVED AND ADVISED THE EROSION WAS CLOSED AND THERE WAS NO INFECTION. THE DEVICE WAS EXPLANTED AND REPLACED SUCCESSFULLY WITH NO PATIENT COMPLICATIONS REPORTED. THE EXPLANTED DEVICE WAS DISPOSED OF AND WILL NOT RETURN FOR ANALYSIS. OUTSIDE OF THE REVISION, NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21212224&pc=LWP | 2124215-2025-03015 | 2024-06-20 04:00 | 2025-01-22 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM (COMPETITOR LEADS) WAS EXPLANTED DUE TO INFECTION. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21212249&pc=LWP | 2124215-2025-03025 | 2024-07-04 04:00 | 2025-01-22 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS PACEMAKER SYSTEM EXPERIENCED AN INFECTION. SUBSEQUENTLY, THE PATIENT UNDERWENT A SURGICAL PROCEDURE, AND THIS SYSTEM WAS EXPLANTED. THIS EXPLANTED DEVICE AND BOTH NON-BOSTON SCIENTIFIC LEADS WERE DISPOSED OF AND WILL NOT RETURN FOR ANALYSIS. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED OUTSIDE OF THE SURGICAL PROCEDURE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21197419&pc=LWP | 2124215-2025-02567 | 2024-12-23 05:00 | 2025-01-21 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection; Pain; Pocket Erosion; Swelling/ Edema | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO POSSIBLE POCKET EROSION. REPORTEDLY THE PATIENT COMPLAINED THAT THE DEVICE WAS STICKING OUT AND MOVING, BUT THE DOCTOR REASSURED THAT THIS WAS TYPICAL AS THE PATIENT WAS THIN. IN ADDITION, THE PATIENT REPORTS THAT THE POCKET LOCATION WAS SWOLLEN AND IT'S PAINFUL TO TOUCH. THERE WAS NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THIS PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21201311&pc=LWP | 2124215-2025-02722 | 2024-10-18 04:00 | 2025-01-21 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Unspecified Infection | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF THE SYSTEM REVISION DUE TO INFECTION. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THE PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21190881&pc=LWP | 2124215-2025-02344 | 2024-12-27 05:00 | 2025-01-20 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO INFECTION. THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THE PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21190888&pc=LWP | 2124215-2025-02350 | 2024-12-26 05:00 | 2025-01-20 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO POCKET INFECTION. THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THIS PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21191770&pc=LWP | 2124215-2025-02395 | 2024-12-26 05:00 | 2025-01-20 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Bacterial Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO METHICILLIN-SUSCEPTIBLE STAPHYLOCOCCUS AUREUS (MSSA) INFECTION. THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THIS PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21193495&pc=LWP | 2124215-2025-02488 | 2024-12-27 05:00 | 2025-01-20 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO INFECTION. THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THIS PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21194244&pc=LWP | 2124215-2025-02513 | 2024-12-21 05:00 | 2025-01-20 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection; Pocket Erosion | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF THE SYSTEM REVISION DUE TO INFECTION AND EROSION. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THE PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21188080&pc=LWP | 2124215-2025-02296 | 2024-12-20 05:00 | 2025-01-19 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Endocarditis | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER WAS EXPLANTED DUE TO ENDOCARDITIS INFECTION. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21182721&pc=LWP | 2124215-2025-02125 | 2025-01-06 05:00 | 2025-01-17 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT UNDERWENT A REVISION SURGERY IN WHICH THIS PACEMAKER WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER HAS BEEN RECEIVED FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21184440&pc=LWP | 2124215-2025-02191 | 2024-12-20 05:00 | 2025-01-17 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection; Inflammation; Pain; Discomfort; Swelling/ Edema | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO INFECTION. REPORTEDLY, THE PATIENT HAD A SYMPTOM OF DISCOMFORT, PAIN, EDEMA, INFLAMMATION AND SWELLING. THERE WAS NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21184446&pc=LWP | 2124215-2025-02187 | 2024-12-18 05:00 | 2025-01-17 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Unspecified Infection; Burning Sensation | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER HAD EXPERIENCED BURNING SENSATION AT THE POCKET SITE. IT WAS REVIEWED THAT THE CONSTANT BURNING SENSATION DOES NOT SOUND ELECTRICAL AS THE DEVICE PACES ONCE PER SECOND. POSSIBLY AN INDICATION OF INFECTION OR ALLERGY. THE TECHNICAL SERVICES (TS) EMAILED THE LIST OF CONTACTING MATERIALS TO THE PATIENT. THERE WAS NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21185783&pc=LWP | 2124215-2025-02202 | 2024-12-19 05:00 | 2025-01-17 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF THE SYSTEM REVISION DUE TO INFECTION. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THE PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21185880&pc=LWP | 2124215-2025-02208 | 2024-12-19 05:00 | 2025-01-17 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF THE SYSTEM REVISION DUE TO INFECTION. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THE PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21169213&pc=LWP | 2124215-2025-01755 | 2024-11-06 05:00 | 2025-01-16 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Unspecified Infection; Pocket Erosion | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF THE SYSTEM REVISION DUE TO INFECTION AND POCKET EROSION. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THE PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21169220&pc=LWP | 2124215-2025-01769 | 2025-01-01 05:00 | 2025-01-16 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Syncope/Fainting | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS PACEMAKER HAD TWO SYNCOPAL EPISODES. TECHNICAL SERVICES (TS) WAS CONSULTED AND REVIEWED STORED DATA, WITH ONLY A RIGHT ATRIAL AUTOMATIC THRESHOLD (RAAT) SUSPENSION AND A RIGHT VENTRICULAR AUTOMATIC THRESHOLD (RVAT) SUSPENSION NOTED AND NO DEVICE ISSUES. NO INFORMATION IS AVAILABLE FROM THE FIELD. THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21174470&pc=LWP | 2124215-2025-01938 | 2024-12-19 05:00 | 2025-01-16 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO INFECTION. THERE WAS NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21158109&pc=LWP | 2124215-2025-01454 | 2024-12-17 05:00 | 2025-01-15 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Convulsion/Seizure | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS IMPLANTABLE PACEMAKER DEVICE EXPERIENCED SEIZURE RESULTING IN MOTOR VEHICULAR ACCIDENT (MVA). BOSTON SCIENTIFIC TECHNICAL SERVICES (TS) REVIEWED THE COLLECTION PERIOD AND APPEARED THAT THE DEVICE APPEARED TO BE FUNCTIONING AS PROGRAMMED. AS OF THIS TIME, THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21161161&pc=LWP | 2124215-2025-01547 | 2024-11-12 05:00 | 2025-01-15 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO INFECTION. A NEW LEADLESS PACEMAKER WAS IMPLANTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21161785&pc=LWP | 2124215-2025-01559 | 2024-12-19 05:00 | 2025-01-15 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Bacterial Infection; Sepsis | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO A SYSTEMIC METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS (MRSA) INFECTION. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21164054&pc=LWP | 2124215-2025-01679 | 2024-12-23 05:00 | 2025-01-15 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF THE SYSTEM REVISION DUE TO INFECTION. THE PATIENT WAS TREATED WITH INTRAVENOUS ANTIBIOTICS. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THE PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21164471&pc=LWP | 2124215-2025-01683 | 2024-09-03 04:00 | 2025-01-15 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Perforation | P150012 | 1 | Event Description: IT WAS REPORTED THAT PHYSICIAN POKED A WHOLE IN PATIENTS' HEART WHEN PUTTING IN THE IMPLANTABLE PACEMAKER. PATIENT STATED HAD TO HAVE A DRAIN FOR SIX WEEKS. BOSTON SCIENTIFIC TECHNICAL SERVICES (TS) REFERRED PATIENT TO PHYSICIAN. THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21145892&pc=LWP | 2124215-2025-01160 | 2024-12-20 05:00 | 2025-01-13 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXPLANTED DUE TO AN UNKNOWN PRODUCT PERFORMANCE ISSUE. THIS PACEMAKER WAS SUCCESSFULLY REPLACED. OTHER THAN THE PROCEDURE, NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER WAS RETURNED TO BOSTON SCIENTIFIC; HOWEVER, FURTHER ANALYSIS HAS NOT BEEN YET PERFORMED.Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXPLANTED DUE TO AN UNKNOWN PRODUCT PERFORMANCE ISSUE. THIS PACEMAKER WAS SUCCESSFULLY REPLACED. OTHER THAN THE PROCEDURE, NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER WAS RETURNED TO BOSTON SCIENTIFIC; HOWEVER, FURTHER ANALYSIS HAS NOT BEEN YET PERFORMED. ADDITIONAL INFORMATION WAS RECEIVED, THIS PACEMAKER WAS ELECTIVELY EXPLANTED AND REPLACED AS PER PHYSICIAN DISCRETION. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER WAS RETURNED TO BOSTON SCIENTIFIC; HOWEVER, FURTHER ANALYSIS HAS NOT BEEN YET COMPLETED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21123676&pc=LWP | 2124215-2025-00613 | 2024-12-12 05:00 | 2025-01-10 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO INFECTION. THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THE PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21123761&pc=LWP | 2124215-2025-00673 | 2024-12-10 05:00 | 2025-01-10 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER WAS EXPLANTED DUE TO UNKNOWN REASON. A NEW DEVICE WAS IMPLANTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21126360&pc=LWP | 2124215-2025-00807 | 2024-12-12 05:00 | 2025-01-10 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO INFECTION. THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THE PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21126946&pc=LWP | 2124215-2025-00836 | 2024-12-13 05:00 | 2025-01-10 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO INFECTION. THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THE PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21108720&pc=LWP | 2124215-2025-00422 | 2024-12-09 05:00 | 2025-01-09 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection; Skin Inflammation/ Irritation | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF THE OF THE SYSTEM REVISION DUE TO INFECTION. REPORTEDLY, THE PATIENT HAD HIVES AND SKIN IRRITATION. NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THE PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21114423&pc=LWP | 2124215-2025-00506 | 2024-12-12 05:00 | 2025-01-09 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI SR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO INFECTION. THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THIS PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21118413&pc=LWP | 2124215-2025-00533 | 2024-12-10 05:00 | 2025-01-09 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Unspecified Infection; Pocket Erosion | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF THE SYSTEM REVISION DUE TO POCKET EROSION AND INFECTION. NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THE PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21118420&pc=LWP | 2124215-2025-00534 | 2024-12-23 05:00 | 2025-01-09 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXPLANTED AND REPLACED WITH A NEW ONE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER HAS NOT BEEN RECEIVED FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21119508&pc=LWP | 2124215-2025-00571 | 2024-12-10 05:00 | 2025-01-09 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF THE SYSTEM REVISION DUE TO INFECTION. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THE PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21102814&pc=LWP | 2124215-2024-83100 | 2024-11-28 05:00 | 2025-01-08 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Unspecified Infection | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER AND LEAD SYSTEM WAS EXPLANTED ABOUT ONE MONTH POST-IMPLANT, DUE TO THE PATIENT'S INFECTION. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THE EXPLANTED PRODUCTS WERE NOT RETURNED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21105138&pc=LWP | 2124215-2025-00270 | 2024-12-09 05:00 | 2025-01-08 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO INFECTION. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21106917&pc=LWP | 2124215-2025-00315 | 2024-12-09 05:00 | 2025-01-08 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF THE SYSTEM REVISION DUE TO INFECTION. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THE PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21107953&pc=LWP | 2124215-2025-00366 | 2024-12-09 05:00 | 2025-01-08 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF THE SYSTEM REVISION DUE TO INFECTION. THE PATIENT WAS TREATED WITH INTRAVENOUS ANTIBIOTICS. NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THE PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21108398&pc=LWP | 2124215-2025-00392 | 2024-12-09 05:00 | 2025-01-08 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF THE SYSTEM REVISION DUE TO INFECTION. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THE PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21081842&pc=LPH | 0002249697-2025-00017 | 2022-07-18 04:00 | 2025-01-06 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE 132 SIZE 2.5 | Foreign Body Reaction | K994366 | 1 | Event Description: IT WAS REPORTED BY COUNSEL THAT CLAIMANT UNDERWENT RIGHT THA ON (B)(6) 2012 AND WAS IMPLANTED WITH AN LFIT V40 FEMORAL HEAD AND ACCOLADE TMZF HIP STEM. HER RIGHT HIP WAS REVISED ON (B)(6) 2022 ALLEGEDLY DUE TO ELEVATED METAL IONS AND MRI FINDINGS CONSISTENT WITH ADVERSE LOCAL TISSUE REACTION. Manufacturer Narrative: IT WAS NOTED THAT THE DEVICE IS NOT AVAILABLE FOR EVALUATION. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, IT WILL BE PROVIDED IN A SUPPLEMENTAL REPORT UPON COMPLETION OF THE INVESTIGATION. REVIEW OF THE DEVICE HISTORY RECORDS INDICATE DEVICES WERE MANUFACTURED AND ACCEPTED INTO FINAL STOCK WITH NO RELEVANT REPORTED DISCREPANCIES. THERE HAVE BEEN NO OTHER SIMILAR EVENTS FOR THE LOT REFERENCED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21064854&pc=LWP | 2124215-2024-82950 | 2024-10-30 04:00 | 2025-01-03 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO WOUND INFECTION. THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THIS PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21055331&pc=LWP | 2124215-2024-82801 | 2024-11-17 05:00 | 2025-01-02 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI SR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO A POCKET INFECTION. FEW WEEKS LATER, A NEW PACING SYSTEM WAS IMPLANTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21045267&pc=LWP | 2124215-2024-82538 | 2024-12-06 05:00 | 2024-12-31 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO INFECTION. THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THIS PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21045390&pc=LWP | 2124215-2024-82554 | 2024-12-06 05:00 | 2024-12-31 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER WAS EXPLANTED DUE TO UNKNOWN ISSUE. A NEW DEVICE WAS SUCCESSFULLY IMPLANTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21045401&pc=LWP | 2124215-2024-82553 | 2024-12-04 05:00 | 2024-12-31 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO INFECTION. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21045418&pc=LWP | 2124215-2024-82560 | 2024-12-09 05:00 | 2024-12-31 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER WAS EXPLANTED DUE TO INFECTION. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21048709&pc=LWP | 2124215-2024-82646 | 2024-12-02 05:00 | 2024-12-31 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Concussion; Syncope/Fainting | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PATIENT WITH THIS PACEMAKER BLACKED OUT. THIS RESULTED IN A FALL AND A CONCUSSION. NO FURTHER INFORMATION WAS PROVIDED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. CURRENTLY, THIS PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21034542&pc=LWP | 2124215-2024-82335 | 2024-12-03 05:00 | 2024-12-30 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO INFECTION. THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THIS PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21035596&pc=LWP | 2124215-2024-82352 | 2024-12-03 05:00 | 2024-12-30 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO INFECTION. THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THIS PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21036716&pc=LWP | 2124215-2024-82357 | 2024-09-24 04:00 | 2024-12-30 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection; Pocket Erosion | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO POCKET EROSION. THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THIS PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21017589&pc=LWP | 2124215-2024-81910 | 2024-10-29 04:00 | 2024-12-26 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER DEVICE WAS SURGICALLY EXPLANTED DUE TO UNKNOWN REASONS. NO DEVICE REPLACEMENT. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER DEVICE WAS SURGICALLY EXPLANTED DUE TO UNKNOWN REASONS. NO DEVICE REPLACEMENT. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS SUPPLEMENTAL CORRECTION REPORT WAS FILED TO CAPTURE THE ADDITIONAL INFORMATION THAT INDICATES THAT THIS DEVICE WAS NEVER IN SERVICE/USED. THIS DOES NOT MEET REPORTABLE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21009643&pc=LWP | 2124215-2024-81585 | 2024-11-29 05:00 | 2024-12-24 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection; Inflammation | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO A POCKET INFECTION. REPORTEDLY, AN INFLAMMATION WAS OBSERVED IN THE POCKET AREA. THERE WAS NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21010010&pc=LWP | 2124215-2024-81685 | 2024-11-27 05:00 | 2024-12-24 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF THE SYSTEM REVISION DUE TO INFECTION. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THE PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21014571&pc=LWP | 2124215-2024-81845 | 2024-11-27 05:00 | 2024-12-24 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Bacterial Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF THE SYSTEM REVISION DUE TO BACTEREMIA. THE PATIENT WAS TREATED WITH INTRAVENOUS ANTIBIOTICS. THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THE PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20996966&pc=LWP | 2124215-2024-81304 | 2024-10-22 04:00 | 2024-12-21 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Unspecified Infection | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF THE SYSTEM REVISION DUE TO INFECTION. REPORTEDLY, THE PATIENT UNDERGONE AN IMPLANT PROCEDURE BUT THE CARDIOPLEGIA WAS NOT WELL FIXED THEN WAS PLACED CORRECTLY WITH GOOD FIXATION. THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THE PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20997007&pc=LWP | 2124215-2024-81322 | 2024-11-25 05:00 | 2024-12-21 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF THE SYSTEM REVISION DUE TO INFECTION. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THE PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20985150&pc=LWP | 2124215-2024-80921 | 2024-11-21 05:00 | 2024-12-20 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF THE SYSTEM REVISION DUE TO INFECTION. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THE PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20990487&pc=LWP | 2124215-2024-81017 | 2024-11-26 05:00 | 2024-12-20 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Endocarditis; Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO ENDOCARDITIS. THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THIS PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20995910&pc=LWP | 2124215-2024-81271 | 2024-11-23 05:00 | 2024-12-20 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection; Pocket Erosion | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER HAD A BULGING OF THE IMPLANTABLE DEVICE AND POSSIBLE EROSION ISSUE. REPORTEDLY, THE PATIENT ADVOCATE NOTICED THE IMPLANT SITE HAS A VISIBLE IMPRINT OF THE PACEMAKER. THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THE PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20978873&pc=LPH | 0002249697-2024-01757 | 2019-12-11 05:00 | 2024-12-19 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE (127 DEG) SIZE 3.5 ACCOLADE (127 DEG) SIZE 3.5 | Pain | K233498 | 1 | Event Description: IT WAS REPORTED BY COUNSEL THAT CLAIMANT UNDERWENT RIGHT THA ON (B)(6) 2011 AND WAS IMPLANTED WITH AN LFIT V40 ANATOMIC FEMORAL HEAD AND ACCOLADE TMZF HIP STEM. SHE WAS REVISED ON (B)(6) 2019 DUE TO A PAINFUL HIP. Manufacturer Narrative: REPORTED EVENT: AN EVENT REGARDING PAIN INVOLVING A ACCOLADE STEM WAS REPORTED. THE EVENT WAS NOT CONFIRMED. METHOD & RESULTS: PRODUCT EVALUATION AND RESULTS: MATERIAL ANALYSIS, VISUAL, FUNCTIONAL AND DIMENSIONAL INSPECTIONS COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED. CLINICIAN REVIEW: NO MEDICAL RECORDS WERE RECEIVED FOR REVIEW WITH A CLINICAL CONSULTANT. PRODUCT HISTORY REVIEW: REVIEW OF THE DEVICE HISTORY RECORDS INDICATE DEVICES WERE MANUFACTURED AND ACCEPTED INTO FINAL STOCK WITH NO RELEVANT REPORTED DISCREPANCIES. COMPLAINT HISTORY REVIEW: THERE HAVE BEEN NO OTHER SIMILAR EVENTS FOR THE LOT REFERENCED. CONCLUSIONS: THE EXACT CAUSE OF THE EVENT COULD NOT BE DETERMINED BECAUSE INSUFFICIENT INFORMATION WAS PROVIDED. FURTHER INFORMATION SUCH AS RETURN OF THE DEVICE, PATHOLOGY REPORTS, PRE- AND POST-OPERATIVE X-RAYS AND THE REVISION OPERATIVE REPORT AS WELL AS PATIENT HISTORY AND FOLLOW-UP NOTES ARE NEEDED TO COMPLETE THE INVESTIGATION FOR DETERMINING ROOT CAUSE. NO FURTHER INVESTIGATION FOR THIS EVENT IS POSSIBLE AT THIS TIME. IF DEVICES AND/OR ADDITIONAL INFORMATION BECOME AVAILABLE TO INDICATE FURTHER EVALUATION IS WARRANTED, THIS RECORD WILL BE REOPENED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20978982&pc=LPH | 0002249697-2024-01766 | 2022-11-17 05:00 | 2024-12-19 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE (127 DEG) SIZE 4.5 ACCOLADE (127 DEG) SIZE 4.5 | Foreign Body Reaction; Pain | K994366 | 1 | Event Description: IT WAS REPORTED BY COUNSEL THAT CLAIMANT UNDERWENT RIGHT THA ON (B)(6) 2011 AND WAS IMPLANTED WITH AN LFIT V40 ANATOMIC FEMORAL HEAD AND ACCOLADE TMZF HIP STEM. HE WAS REVISED ON (B)(6) 2022 ALLEGEDLY DUE TO CHRONIC RIGHT HIP PAIN WITH MILDLY ELEVATED COBALT LEVELS. Manufacturer Narrative: IT WAS NOTED THAT THE DEVICE IS NOT AVAILABLE FOR EVALUATION. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, IT WILL BE PROVIDED IN A SUPPLEMENTAL REPORT UPON COMPLETION OF THE INVESTIGATION. REVIEW OF THE DEVICE HISTORY RECORDS INDICATE DEVICES WERE MANUFACTURED AND ACCEPTED INTO FINAL STOCK WITH NO RELEVANT REPORTED DISCREPANCIES. THERE HAVE BEEN NO OTHER SIMILAR EVENTS FOR THE LOT REFERENCED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20967882&pc=LWP | 2124215-2024-80208 | 2024-11-21 05:00 | 2024-12-18 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Fever; Unspecified Infection | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO INFECTION. REPORTEDLY, THE PATIENT WAS HOSPITALIZED FOR FEVER DUE TO INFECTION AND THE DEVICE WAS REMOVED BECAUSE OF LACK OF IMPROVEMENT. THERE WAS NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE WAS EXPLANTED.Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO INFECTION. REPORTEDLY, THE PATIENT WAS HOSPITALIZED FOR FEVER DUE TO INFECTION AND THE DEVICE WAS REMOVED BECAUSE OF LACK OF IMPROVEMENT. THERE WAS NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE WAS EXPLANTED. Manufacturer Narrative: THIS SUPPLEMENTAL IS BEING FILED TO CAPTURE THE ADDITIONAL INFORMATION OF THE (A) PATIENT INFORMATION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20949286&pc=LWP | 2124215-2024-79701 | 2024-10-09 04:00 | 2024-12-17 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER REACHED ELECTIVE REPLACEMENT INDICATOR (ERI) A MONTH AGO. THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER REACHED ELECTIVE REPLACEMENT INDICATOR (ERI) A MONTH AGO. THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION RECEIVED INDICATES THAT THE DEVICE WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20956373&pc=LWP | 2124215-2024-79952 | 2024-11-20 05:00 | 2024-12-17 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Unspecified Infection; Pocket Erosion | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER HAD A POTENTIAL EROSION ISSUE. REPORTEDLY, THE PATIENT LOST WEIGHT AND THE IMPLANTED DEVICE WAS STICKING OUT. THE HEALTH CARE PROFESSIONAL (HCP) WAS REFERRED TO THE CLINIC. ALL AVAILABLE INFORMATION INDICATED THAT THE PACEMAKER REMAINS IN SERVICE. THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20958614&pc=LWP | 2124215-2024-80086 | 2024-11-26 05:00 | 2024-12-17 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO INFECTION. THERE WAS NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20945884&pc=LWP | 2124215-2024-79533 | 2024-11-19 05:00 | 2024-12-16 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Erosion; Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO EROSION WITH INFECTION. A NEW SYSTEM WAS IMPLANTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20947957&pc=LWP | 2124215-2024-79636 | 2024-12-02 05:00 | 2024-12-16 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Syncope/Fainting | P150012 | 1 | Event Description: IT WAS REPORTED BY THE PATIENT ADVOCATE, THAT THE PATIENT WITH THIS PACEMAKER EXPERIENCED A SYNCOPAL EPISODE. IN ADDITION, THERE WERE CONCERNS ABOUT THE DEVICE PERFORMANCE. TECHNICAL SERVICES (TS) REFERRED THE PATIENT TO THE MEDICAL DOCTOR. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.Event Description: IT WAS REPORTED BY THE PATIENT ADVOCATE, THAT THE PATIENT WITH THIS PACEMAKER EXPERIENCED A SYNCOPAL EPISODE. IN ADDITION, THERE WERE CONCERNS ABOUT THE DEVICE PERFORMANCE. TECHNICAL SERVICES (TS) REFERRED THE PATIENT TO THE MEDICAL DOCTOR. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE. ADDITIONAL INFORMATION WAS RECEIVED, THE FIELD CLINICAL REPRESENTATIVE INDICATED THAT THE DEVICE WAS WORKING AS EXPECTED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20925789&pc=LWP | 2124215-2024-79017 | 2024-11-02 04:00 | 2024-12-13 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI SR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO A POSSIBLE INFECTION. REPORTEDLY, THE PATIENT'S PACER SITE WAS HIT BY A RAKE. AND WAS DEVELOPED INTO A RED SORE, WARM TO TOUCH. THERE WAS NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20929299&pc=LWP | 2124215-2024-79109 | 2024-11-16 05:00 | 2024-12-13 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Unspecified Infection | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO INFECTION. REPORTEDLY, REDNESS WAS OBSERVED AT THE SURGICAL SITE. THE PATIENT WAS TREATED WITH ANTIBIOTICS. THERE WAS NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO INFECTION. REPORTEDLY, REDNESS WAS OBSERVED AT THE SURGICAL SITE. THE PATIENT WAS TREATED WITH ANTIBIOTICS. THERE WAS NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE. Manufacturer Narrative: THIS SUPPLEMENTAL IS BEING FILED TO CORRECT D6A: IMPLANT DATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20929427&pc=LWP | 2124215-2024-79119 | 2024-11-19 05:00 | 2024-12-13 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Unspecified Infection; Sepsis | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF A SYSTEM REVISION DUE TO SEPSIS AND VEGETATION ON LEA. THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THIS PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20930523&pc=LWP | 2124215-2024-79159 | 2024-09-29 04:00 | 2024-12-13 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | Wound Dehiscence | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXPLANTED AND REPLACED, DUE TO WOUND DEHISCENCE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THE EXPLANTED DEVICE WAS NOT EXPECTED TO BE RETURNED.Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXPLANTED AND REPLACED, DUE TO WOUND DEHISCENCE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THE EXPLANTED DEVICE WAS NOT EXPECTED TO BE RETURNED. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THE DEVICE HAD BEEN IMPLANTED FOR THREE DAYS BEFORE IT WAS EXPLANTED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THIS REPORT CORRECTS THE IMPLANT DATE IN D6A. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20920642&pc=LWP | 2124215-2024-78888 | 2024-11-15 05:00 | 2024-12-12 05:00 | Injury | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS PART OF THE SYSTEM REVISION DUE TO INFECTION. THERE WERE NO ADDITIONAL ADVERSE PATIENT EFFECTS REPORTED. THE PACEMAKER WAS EXPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21254054&pc=LWP | 2124215-2025-04131 | 2025-01-02 05:00 | 2025-01-28 05:00 | Injury | Battery Problem | ACCOLADE EL DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER DEVICE HAS DEPLETED FROM 2 YEARS BATTERY REMAINING TO ELECTIVE REPLACEMENT INDICATOR AND DEVICE IS UNDER ADVISORY. THIS DEVICE WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21186186&pc=LWP | 2124215-2025-02269 | 2024-12-19 05:00 | 2025-01-18 05:00 | Injury | Battery Problem | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER LONGEVITY DROPPED FROM 9 YEARS TO ELECTIVE REPLACEMENT INDICATOR (ERI) TO 5.5 YEARS. BOSTON SCIENTIFIC TECHNICAL SERVICES (TS) DISCUSSED THAT PATIENT HAD BEEN ZERO PERCENT RIGHT VENTRICULAR (RV) PACED AND WAS RECENTLY PACED TO 73 PERCENT WITH HIGH OUTPUT. THE DEVICE WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21178609&pc=LWP | 2124215-2025-02041 | 2024-10-23 04:00 | 2025-01-17 05:00 | Injury | Battery Problem | ACCOLADE DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THE LONGEVITY CALCULATION ON THIS IMPLANTABLE PACEMAKER HAS BEEN SITTING AT 3 MONTHS REMAINING FOR A YEAR. DATA WAS SENT FOR ENGINEERING ANALYSIS. ANALYSIS SHOWED THAT THE DEVICE WAS DEPLETING NORMALLY. THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION RECEIVED INDICATES THAT THE DEVICE WAS EXPLANTED AND REPLACED DUE TO BATTERY ADVISORY. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21169215&pc=LWP | 2124215-2025-01766 | 2024-12-20 05:00 | 2025-01-16 05:00 | Injury | Battery Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER DEVICE HAS 4 MONTHS BATTERY REMAINING AND REMAINED THE SAME FOR 2 MONTHS. AS PER CURRENT CHECK, THIS DEVICE HAS LESS THAN 3 MONTHS BATTERY AND IS INCLUDED IN ADVISORY. SUBSEQUENTLY, THIS DEVICE WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21163522&pc=LWP | 2124215-2025-01628 | 2024-12-18 05:00 | 2025-01-15 05:00 | Injury | Battery Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE IMPLANTABLE PACEMAKER DEVICE WAS EXPLANTED DUE TO BATTERY DEPLETION. A NEW DEVICE WAS IMPLANTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21130006&pc=LWP | 2124215-2025-00876 | 2024-10-17 04:00 | 2025-01-10 05:00 | Injury | Battery Problem | ACCOLADE DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER REACHED ELECTIVE REPLACEMENT INDICATOR (ERI). BOSTON SCIENTIFIC TECHNICAL SERVICES (TS) DISCUSSED THAT THIS DEVICE CAN BE PULLED OUT OF ERI IF VOLTAGE AND POWER CONSUMPTION IMPROVE. THIS DEVICE WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21105120&pc=LWP | 2124215-2025-00264 | 2024-10-01 04:00 | 2025-01-08 05:00 | Injury | Battery Problem | ACCOLADE MRI DR | Dyspnea; Unspecified Infection | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER DEVICE HAS A 1 YEAR REMAINING 2 MONTHS AGO. IN ADDITION, THIS PATIENT WAS ADMITTED IN THE EMERGENCY ROOM (ER) FOR URINARY TRACT INFECTION WITH VASCULAR DEMENTIA AND SHORTNESS OF BREATH (SOB) FREQUENTLY. AS OF THIS TIME, THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21009307&pc=LWP | 2124215-2024-81588 | 2024-12-11 05:00 | 2024-12-23 05:00 | Injury | Battery Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE LONGEVITY OF THE BATTERY OF THIS PACEMAKER WAS BEING QUESTIONED. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. THIS DEVICE WAS REPLACED AND IS EXPECTED TO BE RETURNED TO BOSTON SCIENTIFIC FOR ANALYSIS. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21144853&pc=LWP | 2124215-2025-01070 | 2024-12-25 05:00 | 2025-01-13 05:00 | Injury | Battery Problem; Impedance Problem | ACCOLADE MRI DR | Bradycardia; Discomfort | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED IMPEDANCE ISSUES. IT WAS ALSO NOTED THAT THERE WERE CONCERNS ABOUT THIS PACEMAKER'S LONGEVITY. THE PATIENT WITH THIS PACEMAKER EXPERIENCED DISCOMFORT AND BRADYCARDIA. THE PATIENT WAS HOSPITALIZED AND ADMINISTERED MEDICATION TO ACCELERATE THE PATIENT'S HEART RATE. THE DEVICE WAS EXPLANTED AND REPLACED. THE DEVICE IS EXPECTED TO BE RETURNED FOR ANALYSIS. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21021162&pc=LWP | 2124215-2024-82007 | 2024-12-17 05:00 | 2024-12-26 05:00 | Injury | Battery Problem; Interrogation Problem | ACCOLADE MRI DR | Bradycardia | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS UNABLE TO BE INTERROGATED. TROUBLESHOOTING WAS PERFORMED BUT THE ISSUE WAS NOT RESOLVED. TECHNICAL SERVICES (TS) RECOMMENDED PLACING A MAGNET, HOWEVER, NO RESPONSE WAS ELICITED. THE FIELD REPRESENTATIVE STATED THAT THIS DEVICE HAD 2 YEARS OF BATTERY LEFT. A DEVICE CHANGEOUT PROCEDURE WAS PERFORMED AND IT WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. IT IS EXPECTED TO RECEIVE THIS DEVICE FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20926321&pc=LWP | 2124215-2024-79043 | 2024-12-05 05:00 | 2024-12-13 05:00 | Injury | Connection Problem | ACCOLADE MRI EL DR | Loss of consciousness; Syncope/Fainting | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED A SPRING CONTACT PROBLEM WITH A NON-BOSTON SCIENTIFIC RIGHT VENTRICULAR (RV) LEAD, WHICH HAD AN INTEGRITY DEFECT. THIS RESULTED IN THE PATIENT EXPERIENCING SYNCOPE, FAINTING AND LOSS OF CONSCIOUSNESS. SUBSEQUENTLY, THE PATIENT UNDERWENT REVISION SURGERY DURING WHICH BOTH THE PACEMAKER AND THE RV LEAD WERE EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER HAS NOT BEEN RETURNED FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21284893&pc=LWP | 2124215-2025-05056 | 2025-01-17 05:00 | 2025-01-31 05:00 | Injury | Defective Device | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXPLANTED AND REPLACED DUE TO UNKNOWN REASONS. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21263082&pc=LWP | 2124215-2025-04372 | 2025-01-15 05:00 | 2025-01-29 05:00 | Injury | Defective Device | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXPLANTED AND REPLACED DUE TO A NON-SPECIFIC PRODUCT PERFORMANCE ANOMALY. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE HAS NOT BEEN RETURNED.Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXPLANTED AND REPLACED DUE TO A NON-SPECIFIC PRODUCT PERFORMANCE ANOMALY. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE HAS NOT BEEN RETURNED. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT THIS WAS A PROPHYLACTIC EXPLANT. THE DEVICE DID NOT EXHIBIT A PRODUCT PERFORMANCE ANOMALY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21246310&pc=LWP | 2124215-2025-03835 | 2025-01-10 05:00 | 2025-01-27 05:00 | Injury | Defective Device | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXPLANTED AND REPLACED DUE TO UNKNOWN REASONS AT THIS TIME. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. AT THIS TIME, IT IS UNKNOWN IF THIS DEVICE WILL RETURN FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21248880&pc=LWP | 2124215-2025-03941 | 2025-01-10 05:00 | 2025-01-27 05:00 | Injury | Defective Device | ACCOLADE EL DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXPLANTED AND REPLACED DUE TO UNKNOWN PRODUCT PERFORMANCE ISSUE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21198944&pc=LWP | 2124215-2025-02626 | 2024-12-26 05:00 | 2025-01-21 05:00 | Injury | Defective Device | ACCOLADE DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXPLANTED DUE TO A PRODUCT PERFORMANCE ANOMALY. A NEW PACEMAKER WAS SUCCESSFULLY IMPLANTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PRODUCT HAS BEEN RECEIVED BY BOSTON SCIENTIFIC FOR ANALYSIS. Manufacturer Narrative: THE LOCAL AREA SALES REPRESENTATIVE WAS CONTACTED FOR ADDITIONAL INFORMATION. AT THIS TIME, NO FURTHER INFORMATION IS AVAILABLE. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THIS REPORT WILL BE UPDATED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21184459&pc=LWP | 2124215-2025-02177 | 2024-12-19 05:00 | 2025-01-17 05:00 | Injury | Defective Device | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXPLANTED DUE TO AN UNKNOWN PRODUCT PERFORMANCE ISSUE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21152271&pc=LWP | 2124215-2025-01329 | 2024-12-13 05:00 | 2025-01-14 05:00 | Injury | Defective Device | ACCOLADE MRI EL DR | Ventricular Fibrillation | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS IMPLANTABLE PACEMAKER DEVICE EXPERIENCED A VENTRICULAR FIBRILLATION (VF) AND THE ALERT OR NOTIFICATION WAS NOT EVENT UNTIL THE PRESENT TIME. SUBSEQUENTLY, THIS DEVICE WAS EXPLANTED AND UPGRADED TO A TACHY DEVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21119604&pc=LWP | 2124215-2025-00593 | 2024-12-23 05:00 | 2025-01-09 05:00 | Injury | Defective Device | ACCOLADE MRI EL DR | Fall; Dizziness; Syncope/Fainting | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS PACEMAKER SYSTEM EXPERIENCED A SYNCOPAL EPISODE AND FELL. TRANSMISSION ISSUES WERE NOTED ON THE COMMUNICATOR, THEREFORE, THE DEVICE WAS NOT FOUND. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE. FURTHER INFORMATION WAS RECEIVED FROM A PATIENT ADVOCATE RELATED TO THE SYNCOPAL EPISODE AND FALL. THE PATIENT FELT DIZZY, FELL, AND REQUIRED SEVERAL STITCHES ON THEIR HEAD AND WAS TRANSFERRED TO ANOTHER HEALTHCARE FACILITY WHERE THIS PACEMAKER WAS REPROGRAMMED. THE PATIENT HAS REPORTED FEELING BETTER. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE.Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS PACEMAKER SYSTEM EXPERIENCED A SYNCOPAL EPISODE AND FELL. TRANSMISSION ISSUES WERE NOTED ON THE COMMUNICATOR; THEREFORE, THE DEVICE WAS NOT FOUND. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21102816&pc=LWP | 2124215-2024-83082 | 2024-12-18 05:00 | 2025-01-08 05:00 | Injury | Defective Device | ACCOLADE DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS DEVICE WAS EXPLANTED DUE TO AN UNSPECIFIED ADVISORY. A NEW DEVICE WAS SUCCESSFULLY IMPLANTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21108590&pc=LWP | 2124215-2025-00419 | 2024-12-23 05:00 | 2025-01-08 05:00 | Injury | Defective Device | ACCOLADE EL DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THE PACEMAKER WAS EXPLANTED AND REPLACED DUE TO A PRODUCT PERFORMANCE ISSUE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21032390&pc=LWP | 2124215-2024-82286 | 2024-12-17 05:00 | 2024-12-29 05:00 | Injury | Defective Device | ACCOLADE MRI SR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER WAS EXPLANTED DUE TO NON-SPECIFIC PRODUCT PERFORMANCE ISSUE. A NEW DEVICE WAS SUCCESSFULLY IMPLANTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER WAS EXPLANTED DUE TO NON-SPECIFIC PRODUCT PERFORMANCE ISSUE. A NEW DEVICE WAS SUCCESSFULLY IMPLANTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. SUBSEQUENTLY, ADDITIONAL INFORMATION WAS RECEIVED INDICATING DEVICE WAS EXPLANTED DUE TO NORMAL BATTERY DEPLETION. THIS IS NORMAL DEVICE BEHAVIOR AND THERE IS NO MALFUNCTION WITH THE DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED INDICATING DEVICE WAS EXPLANTED DUE TO NORMAL BATTERY DEPLETION. THIS IS NORMAL DEVICE BEHAVIOR AND THERE WAS NO MALFUNCTION WITH THE DEVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21020465&pc=LWP | 2124215-2024-81968 | 2024-12-06 05:00 | 2024-12-26 05:00 | Injury | Defective Device | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXPLANTED AND REPLACED DUE TO UNKNOWN REASONS AT THIS TIME. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20996935&pc=LWP | 2124215-2024-81285 | 2024-11-12 05:00 | 2024-12-21 05:00 | Injury | Defective Device | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER WAS EXPLANTED DUE TO UNKNOWN CAUSE. A NEW DEVICE WAS IMPLANTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20994502&pc=LWP | 2124215-2024-81208 | 2022-08-19 04:00 | 2024-12-20 05:00 | Injury | Defective Device | ACCOLADE MRI SR | Heart Block | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH THIS PACEMAKER WAS HOSPITALIZED ABOUT TWO MONTHS POST-IMPLANT DUE TO AV BLOCK CAUSED BY A SUSPECTED PACEMAKER FAILURE. IT WAS ALSO REPORTED THAT THE PATIENT HAD ANOTHER HOSPITALIZATION LATER (TIMEFRAME NOT KNOWN) DUE TO AN APPARENT PERFORATION OF THE LEFT VENTRICLE. IT IS UNKNOWN WHAT SURGICAL INTERVENTIONS WERE PERFORMED. EVIDENCE SUGGESTS THE DEVICE REMAINS IMPLANTED AND IN SERVICE. AT THIS TIME, THE PATIENT HAS INITIATED LEGAL PROCEEDINGS AGAINST THE IMPLANTING PHYSICIAN, THE HOSPITAL, AND BOSTON SCIENTIFIC. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20960528&pc=LWP | 2124215-2024-80128 | 2024-12-09 05:00 | 2024-12-18 05:00 | Injury | Defective Device | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER WAS EXPLANTED DUE TO NON-SPECIFIC PRODUCT PERFORMANCE ISSUE. A NEW DEVICE WAS SUCCESSFULLY IMPLANTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20970702&pc=LWP | 2124215-2024-80433 | 2024-12-02 05:00 | 2024-12-18 05:00 | Injury | Defective Device | ACCOLADE EL DR | No Clinical Signs, Symptoms or Conditions | N970003/S167 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED AN UNSPECIFIED PRODUCT PERFORMANCE ISSUE. SURGICAL INTERVENTION WAS UNDERTAKEN. THE DEVICE WAS EXPLANTED AND REPLACED WITH A NON-BOSTON SCIENTIFIC DEVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THE DEVICE IS IN HOSPITAL POSSESSION. AT THIS TIME, THE PRODUCT HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED, ANALYSIS WILL BE PERFORMED AND THIS REPORT WOULD BE UPDATED AT THAT TIME.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20922527&pc=LWP | 2124215-2024-78925 | 2024-12-03 05:00 | 2024-12-13 05:00 | Injury | Defective Device | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER WAS EXPLANTED DUE TO NON-SPECIFIC PRODUCT PERFORMANCE ISSUE. A NEW DEVICE WAS SUCCESSFULLY IMPLANTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER WAS EXPLANTED DUE TO NON-SPECIFIC PRODUCT PERFORMANCE ISSUE. A NEW DEVICE WAS SUCCESSFULLY IMPLANTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. SUBSEQUENTLY, ADDITIONAL INFORMATION WAS RECEIVED INDICATING DEVICE WAS EXPLANTED DUE TO NORMAL BATTERY DEPLETION. THERE WAS NO MALFUNCTIONS WITH THE DEVICE. Manufacturer Narrative: ADDITIONAL INFORMATION WAS RECEIVED INDICATING DEVICE WAS EXPLANTED DUE TO NORMAL BATTERY DEPLETION. THERE WAS NO MALFUNCTION WITH THE DEVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20920666&pc=LWP | 2124215-2024-78792 | 2024-11-22 05:00 | 2024-12-12 05:00 | Injury | Defective Device | ACCOLADE MRI DR | Fall; Syncope/Fainting | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS FELL ON HIS HOUSE DUE TO A POSSIBLE SYNCOPAL EPISODE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21114428&pc=LWP | 2124215-2025-00516 | 2024-12-23 05:00 | 2025-01-09 05:00 | Injury | Defective Device; Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXPLANTED AND REPLACED DUE TO A PRODUCT PERFORMANCE ISSUE. OTHER THAN THE SURGICAL PROCEDURE, NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER HAS BEEN RETURNED FOR ANALYSIS.Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXPLANTED AND REPLACED DUE TO A PRODUCT PERFORMANCE ISSUE. OTHER THAN THE SURGICAL PROCEDURE, NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER HAS BEEN RETURNED FOR ANALYSIS. ADDITIONAL INFORMATION WAS RECEIVED THAT THIS PACEMAKER WAS EXPLANTED PRIOR TO REACHING ELECTIVE REPLACEMENT INDICATOR (ERI) SINCE THE PATIENT IS PACEMAKER DEPENDENT. THEREFORE, NO DEVICE MALFUNCTION OCCURRED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER HAS BEEN RETURNED FOR ANALYSIS. Manufacturer Narrative: CORRECTION: CODING H6.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20921803&pc=LWP | 2124215-2024-78833 | 2024-11-22 05:00 | 2024-12-13 05:00 | Injury | Defective Device; Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI SR | Unspecified Infection; No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXPLANTED SHORTLY AFTER IMPLANT DUDE TO UNKNOWN REASONS. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXPLANTED SHORTLY AFTER IMPLANT DUDE TO UNKNOWN REASONS. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION WAS RECEIVED, THIS PACEMAKER SYSTEM WAS EXPLANTED DUE TO A POCKET INFECTION. THE PATIENT WAS IMPLANTED WITH A PACEMAKER FROM A DIFFERENT MANUFACTURER. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21080571&pc=LPH | 0002249697-2025-00005 | 2023-06-15 04:00 | 2025-01-06 05:00 | Injury | Detachment of Device or Device Component | ACCOLADE (127 DEG) SIZE 4.5 ACCOLADE (127 DEG) SIZE 4.5 | Insufficient Information | K994366 | 1 | Event Description: IT WAS REPORTED BY COUNSEL THAT CLAIMANT UNDERWENT RIGHT THA ON (B)(6) 2013 AND WAS IMPLANTED WITH AN LFIT ANATOMIC V40 FEMORAL HEAD AND ACCOLADE TMZF HIP STEM. HIS RIGHT HIP WAS REVISED ON (B)(6) 2023 ALLEGEDLY DUE TO HEAD DISASSOCIATION. Manufacturer Narrative: IT WAS NOTED THAT THE DEVICE IS NOT AVAILABLE FOR EVALUATION. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, IT WILL BE PROVIDED IN A SUPPLEMENTAL REPORT UPON COMPLETION OF THE INVESTIGATION. REVIEW OF THE DEVICE HISTORY RECORDS INDICATE DEVICES WERE MANUFACTURED AND ACCEPTED INTO FINAL STOCK WITH NO RELEVANT REPORTED DISCREPANCIES. THERE HAVE BEEN NO OTHER SIMILAR EVENTS FOR THE LOT REFERENCED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20935736&pc=LWP | 2124215-2024-79266 | 2024-11-15 05:00 | 2024-12-15 05:00 | Injury | Device Contaminated at the User Facility | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER WAS CONTAMINATED ACCIDENTALLY DURING THE PROCEDURE INVOLVING LEADS EXPLANTS. AS A RESULT, THE DEVICE WAS EXPLANTED AND REPLACED WITH A NEW PACEMAKER. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20920454&pc=LPH | 0002249697-2024-01716 | 2023-05-12 04:00 | 2024-12-12 05:00 | Injury | Device Dislodged or Dislocated | ACCOLADE (127 DEG) SIZE 5.5 ACCOLADE (127 DEG) SIZE 5.5 | Insufficient Information | K994366 | 1 | Event Description: IT WAS REPORTED BY COUNSEL THAT CLAIMANT UNDERWENT LEFT THA ON (B)(6) 2011 AND WAS IMPLANTED WITH AN LFIT ANATOMIC V40 FEMORAL HEAD AND ACCOLADE TMZF HIP STEM. HIS LEFT HIP WAS REVISED ON (B)(6) 2023 ALLEGEDLY DUE TO HEAD DISASSOCIATION. Manufacturer Narrative: IT WAS NOTED THAT THE DEVICE IS NOT AVAILABLE FOR EVALUATION. IF ADDITIONAL INFORMATION IS RECEIVED, IT WILL BE PROVIDED IN A SUPPLEMENTAL REPORT UPON COMPLETION OF THE INVESTIGATION. REVIEW OF THE DEVICE HISTORY RECORDS INDICATE DEVICES WERE MANUFACTURED AND ACCEPTED INTO FINAL STOCK WITH NO REPORTED DISCREPANCIES. THERE HAVE BEEN NO OTHER EVENTS FOR THE LOT REFERENCED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21200789&pc=LWP | 2124215-2025-02699 | 2025-01-05 05:00 | 2025-01-21 05:00 | Injury | Failure to Capture | ACCOLADE MRI EL DR | Asystole | P150012 | 1 | Event Description: IT WAS REPORTED THAT PACEMAKER EXHIBITED POTENTIAL LOSS OF CAPTURE, AND THE PATIENT'S HEARTRATE WAS IN THE 30 TO 40 BEATS-PER-MINUTE RANGE. THE PATIENT WAS ADMITTED TO THE HOSPITAL. THE PACEMAKER SYSTEM REMAINS IN SERVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21191589&pc=LWP | 2124215-2025-02372 | 2025-01-04 05:00 | 2025-01-20 05:00 | Injury | Failure to Capture | ACCOLADE MRI DR | Bradycardia; Syncope/Fainting | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS PACEMAKER HAD A SYNCOPE EPISODE DUE TO LOSS OF CAPTURE (LOC). THE IMPEDANCE MEASUREMENTS WERE IN RANGE, HOWEVER, THE PACEMAKER SYSTEM WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER HAS NOT BEEN RETURNED TO BOSTON SCIENTIFIC FOR FURTHER ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21164061&pc=LWP | 2124215-2025-01670 | 2025-01-01 05:00 | 2025-01-15 05:00 | Injury | Failure to Capture | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXPLANTED DUE TO IT EXHIBITING LOSS OF CAPTURE (LOC) FOR ALL ITS CHANNELS. A NEW DEVICE WAS IMPLANTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21231284&pc=LWP | 2124215-2025-03458 | 2025-01-04 05:00 | 2025-01-24 05:00 | Injury | Failure to Capture; High Capture Threshold | ACCOLADE MRI DR | Chest Pain; Cardiac Tamponade; Cardiac Perforation; Pericardial Effusion | P150012 | 1 | Event Description: IT WAS REPORTED THAT PACEMAKER SYSTEM EXHIBITED HIGH PACING THRESHOLDS ON THE RIGHT VENTRICULAR (RV) LEAD, IN ADDITION INTERMITTENT LOSS OF CAPTURE (LOC) WAS ALSO OBSERVED. THE PATIENT FELT CHEST PAIN. AS A RESULT, THE PATIENT UNDERWENT A REVISION PROCEDURE, AND IT WAS FOUND THAT THE PATIENT HAD A CARDIAC TAMPONADE WITH PERICARDIAL EFFUSION DUE TO A PERFORATION OF THE HEART WALL. A PERICARDIAL DRAINED WAS PERFORMED. THE PACEMAKER SYSTEM WAS EXPLANTED AND SUCCESSFULLY REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT PACEMAKER SYSTEM EXHIBITED HIGH PACING THRESHOLDS ON THE RIGHT VENTRICULAR (RV) LEAD, IN ADDITION INTERMITTENT LOSS OF CAPTURE (LOC) WAS ALSO OBSERVED. THE PATIENT FELT CHEST PAIN. AS A RESULT, THE PATIENT UNDERWENT A REVISION PROCEDURE, AND IT WAS FOUND THAT THE PATIENT HAD A CARDIAC TAMPONADE WITH PERICARDIAL EFFUSION DUE TO A PERFORATION OF THE HEART WALL. A PERICARDIAL DRAINED WAS PERFORMED. THE PACEMAKER SYSTEM WAS EXPLANTED AND SUCCESSFULLY REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER IS NOT EXPECTED TO BE RETURNED TO BOSTON SCIENTIFIC SINCE IT WAS RETAINED BY THE EXPLANTING HOSPITAL.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20918537&pc=LWP | 2124215-2024-78840 | 2024-11-15 05:00 | 2024-12-12 05:00 | Injury | Failure to Capture; High Capture Threshold | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER RECORDED A LOSS OF CAPTURE (LOC) ON THE RIGHT VENTRICULAR (RV) LEAD. TECHNICAL SERVICES (TS) WAS CONSULTED AND ADVISED THE LOC APPEARS TO BE INTERMITTENT. THE POLARITY WAS CHANGED FROM BIPOLAR TO UNIPOLAR AND IMPEDANCES ARE NORMAL. THE RIGHT VENTRICULAR AUTOMATIC THRESHOLD (RVAT) TEST WAS TURNED OFF IN THE PAST DUE TO ELEVATED THRESHOLDS. THERE IS NO EVIDENCE OF ASYSTOLE DUE TO THE LOC. TS RECOMMENDED THE PATIENT TO BE EVALUATED IN-CLINIC FOR FURTHER EVALUATION. AT THIS TIME, THE DEVICE AND LEAD REMAIN IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION RECEIVED ADVISED THE PATIENT WAS ADMITTED TO THE HOSPITAL UNTIL THE LEAD REVISION PROCEDURE. AT THIS TIME, THE DEVICE AND LEAD REMAIN IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER RECORDED A LOSS OF CAPTURE (LOC) ON THE RIGHT VENTRICULAR (RV) LEAD. TECHNICAL SERVICES (TS) WAS CONSULTED AND ADVISED THE LOC APPEARS TO BE INTERMITTENT. THE POLARITY WAS CHANGED FROM BIPOLAR TO UNIPOLAR AND IMPEDANCES ARE NORMAL. THE RIGHT VENTRICULAR AUTOMATIC THRESHOLD (RVAT) TEST WAS TURNED OFF IN THE PAST DUE TO ELEVATED THRESHOLDS. THERE IS NO EVIDENCE OF ASYSTOLE DUE TO THE LOC. TS RECOMMENDED THE PATIENT TO BE EVALUATED IN-CLINIC FOR FURTHER EVALUATION. AT THIS TIME, THE DEVICE AND LEAD REMAIN IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION RECEIVED ADVISED THE PATIENT WAS ADMITTED TO THE HOSPITAL UNTIL THE LEAD REVISION PROCEDURE. AT THIS TIME, THE DEVICE AND LEAD REMAIN IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. RECEIVED ADDITIONAL INFORMATION THE DEVICE AND LEAD WERE EXPLANTED AND REPLACED SUCCESSFULLY. OUTSIDE OF THE REVISION, NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21209709&pc=LWP | 2124215-2025-02940 | 2022-09-18 04:00 | 2025-01-22 05:00 | Injury | Failure to Capture; High impedance | ACCOLADE MRI EL DR | Syncope/Fainting | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PATIENT EXPERIENCED SYNCOPAL EPISODES IN THE HOSPITAL. UPON REVIEW, IT WAS NOTED THE LOSS OF CAPTURE ON THE RIGHT VENTRICULAR (RV) CHANNEL, WHICH LED TO PACING INHIBITION GREATER THAN TWO SECONDS OF ASYSTOLE. IT WAS BELIEVED TO BE CAUSED BY A LEAD PERFORATION. ADDITIONALLY, HIGH PACING IMPEDANCE MEASUREMENTS WERE ALSO NOTED. THE PATIENT WAS ADMITTED TO THE HOSPITAL AND A REVISION PROCEDURE WAS PERFORMED. THE RV LEAD WAS REPOSITIONED AND REMAINS IN SERVICE. THEN, A REVISION PROCEDURE WAS PERFORMED, THE RV LEAD AND PACEMAKER WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21044952&pc=LWP | 2124215-2024-82493 | 2024-12-12 05:00 | 2024-12-30 05:00 | Injury | Failure to Capture; Low impedance | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM EXHIBITED INTERMITTENT LOSS OF CAPTURE (LOC) ON THE RIGHT VENTRICULAR (RV) LEAD. ADDITIONALLY, IMPEDANCE MEASUREMENTS WERE NOTED TO VARY FROM 400 TO 300 OHMS. TECHNICAL SERVICES (TS) WAS CONSULTED AND A CHEST X RAY WAS RECOMMENDED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE. FURTHER INFORMATION WAS RECEIVED THAT THE PATIENT WAS HOSPITALIZED AND A NON BOSTON SCIENTIFIC LEADLESS DEVICE WAS IMPLANTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. AT THIS TIME, IT IS UNKNOWN IF THIS DEVICE WILL RETURN FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21277624&pc=LWP | 2124215-2025-04892 | 2025-01-13 05:00 | 2025-01-30 05:00 | Injury | Failure to Capture; Low impedance; Battery Problem; High Capture Threshold | ACCOLADE MRI DR | Bradycardia; Device Overstimulation of Tissue | P150012 | 1 | Event Description: IT WAS REPORTED THAT ABOUT TWO MONTHS POST IMPLANT, THE PATIENT WITH THIS PACEMAKER SYSTEM VISITED THE EMERGENCY ROOM (ER). DURING INTERROGATION THE DEVICE WAS FOUND EXHIBIT RIGHT VENTRICULAR (RV) LOSS OF CAPTURE (LOC) WITH UNDERLYING ESCAPE RHYTHM OF 38 BEATS PER MINUTE (BPM), A DROP IN BATTERY LONGEVITY FROM 8.5YRS TO 1.5 YRS IN JUST TWO MONTHS AND AN ALERT THAT INDICATED THAT THE RIGHT VENTRICULAR AUTOMATIC THRESHOLD (RVAT)RIGHT VENTRICULAR AUTO THRESHOLD (RVAT) TEST WAS IN SUSPENSION. THE HEALTH CARE PROFESSIONAL (HCP) CALLED TECHNICAL SERVICES (TS) FOR REVIEW OF THE PRESENTING ELECTROGRAM (EGM) AND REQUESTED DISK DATA ANALYSIS. THE HCP NOTED THAT THE RV THRESHOLD OUTPUTS WERE INCREASED TO MAX OUTPUTS AND CAPTURE WAS CONFIRMED AT 60BPMS AND THE PATIENT WAS ADMITTED FOR FURTHER OBSERVATIONS. IT WAS NOTED THAT THE RV LEAD IS A NON-BOSTON SCIENTIFIC LEAD. UPON REVIEW, TS NOTED THAT THE RVAT WAS SUSPENDED DUE TO HIGH PACING THRESHOLDS ON THE RV LEAD. TS ENGINEERING ANALYZED THE DEVICE DATA PROVIDED AND NOTED A DECREASE IN RV PACING THRESHOLDS AND IMPEDANCES AND FOUND THE DEVICE POWER CONSUMPTION TO BE STABLE. TS ENGINEERING POSTULATED THE POSSIBLE CHANGE IN BATTERY LONGEVITY ESTIMATES MAY HAVE BEEN DUE TO THE CHANGE IN LEAD OUTPUTS AND RECOMMENDED FOR THE HCP TO REINTERROGATE THE DEVICE TO CONFIRM THE BATTERY LONGEVITY. THE FOLLOWING DAY, THE HCP CALLED TS BACK REPORTING THE PATIENT HAD EXPERIENCED MUSCLE STIMULATION OVERNIGHT. THE HCP ATTEMPTED TO RECREATE THE STIMULATION, HOWEVER, WAS UNSUCCESSFUL. THE HCP ALSO REINTERROGATED THE DEVICE AND THE BATTERY LONGEVITY ESTIMATES WERE BACK TO READ 8.5 YRS, THE RV THRESHOLDS WERE STILL HIGH. THE HCP REPORTED PLANS TO SCHEDULE A REVISION PROCEDURE. TS RECOMMENDED FOR THE HCP TO PERFORM A PACING SYSTEM ANALYZER (PSA) TEST DURING THE REVISION PROCEDURE TO CONFIRM RV LEAD ISSUE. SUBSEQUENTLY, A REVISION PROCEDURE WAS PERFORMED, THE PSA TEST WAS COMPLETED AND SHOWED HIGH PACING THRESHOLDS ON THE NON-BOSTON SCIENTIFIC RV LEAD. THE PHYSICIAN OPTED TO REPLACE THE ENTIRE PACEMAKER SYSTEM. THE PACEMAKER DEVICE WAS EXPLANTED AND REPLACED WITH A NEW DEVICE. WHILE THE NON-BOSTON SCIENTIFIC RIGHT ATRIAL (RA) AND RIGHT VENTRICULAR (RV) LEADS WERE SURGICALLY ABANDONED, AND THE NON-BOSTON SCIENTIFIC RV LEAD WAS REPLACED WITH A NEW LEAD. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THE PACEMAKER IS EXPECTED TO RETURN TO FACILITY.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21211592&pc=LWP | 2124215-2025-02978 | 2025-01-08 05:00 | 2025-01-22 05:00 | Injury | Failure to Capture; Pacing Problem | ACCOLADE MRI EL DR | Asystole | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE PACEMAKER OBSERVED POSSIBLE LOSS OF CAPTURE ON THE ELECTROGRAM (EGM). THE PATIENT'S HEARTRATE WAS IN THE 30 BEATS-PER-MINUTE RANGE AND THEY WERE ADMITTED TO THE HOSPITAL. PACEMAKER MEDIATED TACHYCARDIA (PMT) EPISODES WERE ALSO NOTED. THE PACEMAKER SYSTEM REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21242099&pc=LWP | 2124215-2025-03743 | 2024-12-13 05:00 | 2025-01-27 05:00 | Injury | Failure to Read Input Signal | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT AN ALERT WAS DETECTED FOR RIGHT ATRIAL AUTOMATIC THRESHOLD (RAAT) SUSPENSION. TECHNICAL SERVICES REVIEWED DEVICE DATA AND DETERMINED THE RAAT FAILED TO MEASURE THE THRESHOLD DUE TO LOW EVOKED RESPONSE (ER) BUT A FEW DAYS LATER IT WAS SUCCESSFUL. THIS PATIENT IS BEING MONITORED REMOTELY. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE WAS EXPLANTED AND REPLACED.Event Description: IT WAS REPORTED THAT AN ALERT WAS DETECTED FOR RIGHT ATRIAL AUTOMATIC THRESHOLD (RAAT) SUSPENSION. TECHNICAL SERVICES REVIEWED DEVICE DATA AND DETERMINED THE RAAT FAILED TO MEASURE THE THRESHOLD DUE TO LOW EVOKED RESPONSE (ER) BUT A FEW DAYS LATER IT WAS SUCCESSFUL. THIS PATIENT IS BEING MONITORED REMOTELY. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE WAS EXPLANTED AND REPLACED. Manufacturer Narrative: CORRECTED FIELDS: BSC AWARE DATE G3.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21089468&pc=LPH | 0002249697-2025-00026 | 2024-12-10 05:00 | 2025-01-07 05:00 | Injury | Fracture | ACCOLADE PLUS TMZF HIP STEM #4 | Insufficient Information | K994366 | 1 | Event Description: PATIENT WAS REVISED DUE TO A FRACTURE OF AN IMPLANTED HIP STEM (PRIMARY SURGERY). Manufacturer Narrative: IT WAS NOTED THAT THE DEVICE IS NOT AVAILABLE FOR EVALUATION. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, IT WILL BE PROVIDED IN A SUPPLEMENTAL REPORT UPON COMPLETION OF THE INVESTIGATION. REVIEW OF THE DEVICE HISTORY RECORDS INDICATE DEVICES WERE MANUFACTURED AND ACCEPTED INTO FINAL STOCK WITH NO RELEVANT REPORTED DISCREPANCIES. THERE HAVE BEEN NO OTHER SIMILAR EVENTS FOR THE LOT REFERENCED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21063351&pc=LWP | 2124215-2025-00092 | 2023-02-14 05:00 | 2025-01-02 05:00 | Injury | High impedance | ACCOLADE EL DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT, OVER THE LAST YEAR, THIS RIGHT VENTRICULAR (RV) LEAD EXHIBITED A GRADUAL INCREASE IN PACE IMPEDANCE MEASUREMENTS. THERE WAS ALSO AN INCREASE IN THE PACING THRESHOLD MEASUREMENTS. WHICH WAS SUSPECTED LEAD BROKEN. SUBSEQUENTLY, A REVISION PROCEDURE WAS PERFORMED AND THE RV LEAD WAS SURGICALLY ABANDONED. WHILE THE PACEMAKER WAS EXPLANTED. BOTH PRODUCTS WERE SUCCESSFULLY REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21123709&pc=LWP | 2124215-2025-00592 | 2024-12-20 05:00 | 2025-01-10 05:00 | Injury | High impedance; Defective Device | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXPLANTED DUE TO AN UNKNOWN PRODUCT PERFORMANCE ISSUE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. AT THIS TIME, IT IS UNKNOWN IF THIS DEVICE WILL RETURN FOR ANALYSIS.Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXPLANTED DUE TO AN UNKNOWN PRODUCT PERFORMANCE ISSUE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. FURTHER INFORMATION WAS RECEIVED THAT THE PHYSICIAN COULD NOT RULE OUT THAT HIGH OUT OF RANGE PACE IMPEDANCES WERE ONLY LEAD RELATED AND DECIDED TO EXPLANT THIS DEVICE ALONG WITH THE RIGHT VENTRICULAR (RV) LEAD. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. IT IS EXPECTED TO RECEIVE THIS DEVICE FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21233148&pc=LWP | 2124215-2025-03526 | 2024-12-31 05:00 | 2025-01-24 05:00 | Injury | High impedance; Over-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS SUCCESSFULLY EXPLANTED AND REPLACED DUE TO OVERSENSING ON THE RIGHT ATRIAL (RA) CHANNEL. NOISE WAS NOTED ON THE RA CHANNEL THAT WAS OVERSENSED AND THIS CAUSED THAT THE MINUTE VENTILATION (MV) SENSOR WAS AUTOMATICALLY DISABLED. ADDITIONALLY, HIGH OUT OF RANGE IMPEDANCE MEASUREMENTS WERE NOTED ON THE RA LEAD. SUBSEQUENTLY, THIS PACEMAKER SYSTEM WAS REPLACED WITH A NON-BOSTON SCIENTIFIC SYSTEM. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED OUTSIDE THE REVISION PROCEDURE. THIS PRODUCT HAS NOT RETURNED FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21212250&pc=LWP | 2124215-2025-03020 | 2025-01-07 05:00 | 2025-01-22 05:00 | Injury | Impedance Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED IMPEDANCE ISSUES. SUBSEQUENTLY, THE PATIENT UNDERWENT A REVISION SURGERY IN WHICH THIS DEVICE WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21284892&pc=LWP | 2124215-2025-05061 | 2025-01-17 05:00 | 2025-01-31 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS FOUND TO BE OPERATING IN SAFETY MODE. SUBSEQUENTLY, THIS PACEMAKER WAS EXPLANTED AND REPLACED. OTHER THAN THE SURGICAL PROCEDURE, NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER HAS NOT BEEN RETURNED FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21289001&pc=LWP | 2124215-2025-05206 | 2025-01-15 05:00 | 2025-01-31 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE EL DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER DEVICE WAS EXPLANTED AFTER BEING FOUND TO BE IN SAFETY MODE. THE DEVICE WAS SUCCESSFULLY REMOVED AND REPLACED WITH A NEW DEVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21289301&pc=LWP | 2124215-2025-05215 | 2025-01-15 05:00 | 2025-01-31 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | Unspecified Heart Problem | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS PACEMAKER VISITED THE EMERGENCY ROOM (ER) DUE TO PALPITATION SYMPTOMS DURING UNIPOLAR PACING. THE DEVICE WAS INTERROGATED AND FOUND TO BE OPERATING IN SAFETY MODE. SUBSEQUENTLY, A DEVICE REPLACEMENT PROCEDURE WAS PERFORMED. THE PACEMAKER WAS EXPLANTED AND REPLACED WITH A NEW DEVICE SUCCESSFULLY. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS PACEMAKER VISITED THE EMERGENCY ROOM (ER) DUE TO PALPITATION SYMPTOMS DURING UNIPOLAR PACING. THE DEVICE WAS INTERROGATED AND FOUND TO BE OPERATING IN SAFETY MODE. SUBSEQUENTLY, A DEVICE REPLACEMENT PROCEDURE WAS PERFORMED. THE PACEMAKER WAS EXPLANTED AND REPLACED WITH A NEW DEVICE SUCCESSFULLY. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: CORRECTION: E1: INITIAL REPORTER LAST NAME.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21271131&pc=LWP | 2124215-2025-04687 | 2025-01-16 05:00 | 2025-01-30 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THE PACEMAKER WAS FOUND TO BE OPERATING IN SAFETY MODE. SUBSEQUENTLY, THE PATIENT UNDERWENT A REVISION PROCEDURE AND THE PACEMAKER WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21264653&pc=LWP | 2124215-2025-04406 | 2025-01-15 05:00 | 2025-01-29 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS FOUND TO BE OPERATING IN SAFETY MODE. TECHNICAL SERVICES (TS) ADVISED URGENT DEVICE REPLACEMENT. SUBSEQUENTLY, THIS PACEMAKER WAS EXPLANTED AND REPLACED WITH A NEW ONE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PRODUCT HAS BEEN RECEIVED BY BOSTON SCIENTIFIC FOR FURTHER INVESTIGATION. Manufacturer Narrative: THIS PRODUCT HAS BEEN RECEIVED BY BOSTON SCIENTIFIC AND FULL INVESTIGATION WILL BE PERFORMED. ONCE THE INVESTIGATION IS COMPLETED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED TO PROVIDE ANY NEW INFORMATION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21242865&pc=LWP | 2124215-2024-78205 | 2024-11-20 05:00 | 2025-01-27 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER REVERTED TO SAFETY MODE. THE PATIENT WAS DEPENDENT. THIS PACEMAKER WAS EXPLANTED AND SUCCESSFULLY REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER HAS NOT RETURNED TO BOSTON SCIENTIFIC FOR FURTHER ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21246353&pc=LWP | 2124215-2025-03843 | 2024-12-30 05:00 | 2025-01-27 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER HAD ENTERED SAFETY MODE AND WAS EXPLANTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PRODUCT HAS RETURNED FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21249098&pc=LWP | 2124215-2025-04006 | 2025-01-07 05:00 | 2025-01-27 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER DEVICE WAS EXPLANTED AFTER BEING FOUND TO BE IN SAFETY MODE. THE DEVICE WAS SUCCESSFULLY REMOVED AND REPLACED WITH A NEW DEVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THE DEVICE WAS RETURNED TO FACILITY FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21230038&pc=LWP | 2124215-2025-03423 | 2025-01-06 05:00 | 2025-01-24 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | Syncope/Fainting | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS PACEMAKER HAD A SYNCOPAL EPISODE. WHEN EXAMINED, IT WAS FOUND THE PACEMAKER HAD ENTERED SAFETY MODE. SUBSEQUENTLY, IT WAS EXPLANTED. A NEW DEVICE WAS IMPLANTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21231306&pc=LWP | 2124215-2025-03469 | 2025-01-10 05:00 | 2025-01-24 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS FOUND TO BE OPERATING IN SAFETY MODE. TECHNICAL SERVICES (TS) WAS CONSULTED AND RECOMMENDED DEVICE REPLACEMENT. SUBSEQUENTLY, THE PATIENT UNDERWENT A REVISION SURGERY IN WHICH THIS PACEMAKER WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER HAS NOT BEEN RECEIVED FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21211617&pc=LWP | 2124215-2025-02982 | 2024-12-23 05:00 | 2025-01-22 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER REVERTED TO SAFETY MODE. THE DEVICE WAS EXPLANTED AND REPLACED. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21179766&pc=LWP | 2124215-2025-02061 | 2025-01-03 05:00 | 2025-01-17 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI SR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER REVERTED TO SAFETY MODE. SUBSEQUENTLY, IT WAS EXPLANTED AND SUCCESSFULLY REPLACED. OTHER THAN PROCEDURE, NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER IS NOT EXPECTED TO BE RETURNED TO BOSTON SCIENTIFIC, SINCE WAS RETAINED BY THE EXPLANTING HOSPITAL.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21182743&pc=LWP | 2124215-2025-02124 | 2025-01-03 05:00 | 2025-01-17 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS FOUND TO BE OPERATING IN SAFETY MODE. DEVICE REPLACEMENT WAS RECOMMENDED. SUBSEQUENTLY, A REVISION PROCEDURE WAS PERFORMED AND THE PACEMAKER WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21170004&pc=LWP | 2124215-2025-01788 | 2025-01-02 05:00 | 2025-01-16 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE DR | Insufficient Information | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS FOUND TO BE OPERATING IN SAFETY MODE. ADDITIONALLY, THE PATIENT EXPERIENCED UNSPECIFIED SYMPTOMS RELATED. TECHNICAL SERVICES (TS) WAS CONSULTED AND RECOMMENDED DEVICE REPLACEMENT. SUBSEQUENTLY, THE PATIENT UNDERWENT A REVISION SURGERY IN WHICH THIS PACEMAKER WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER HAS NOT BEEN RECEIVED FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21173299&pc=LWP | 2124215-2025-01854 | 2025-01-13 05:00 | 2025-01-16 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | Syncope/Fainting | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT DURING A FOLLOW UP IN (B)(6) 2024 THE PATIENT'S BATTERY SHOWED TWO YEARS REMAINING. MOST RECENTLY THE PATIENT PRESENTED TO THE EMERGENCY ROOM WITH SYNCOPE AND UPON DEVICE INTERROGATION IT WAS NOTED THAT THE DEVICE WAS IN SAFETY MODE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THE DEVICE REMAINS IMPLANTED.Event Description: IT WAS REPORTED THAT DURING A FOLLOW UP IN (B)(6) 2024 THE PATIENT'S BATTERY SHOWED TWO YEARS REMAINING. MOST RECENTLY THE PATIENT PRESENTED TO THE EMERGENCY ROOM WITH SYNCOPE AND UPON DEVICE INTERROGATION IT WAS NOTED THAT THE DEVICE WAS IN SAFETY MODE. THE DEVICE WAS EXPLANTED AND EXPECTED TO BE RETURNED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS REPORT IS BEING FILED TO CORRECT THE EXPLANT DATE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21158651&pc=LWP | 2124215-2025-01477 | 2024-12-30 05:00 | 2025-01-15 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER REVERTED TO SAFETY MODE. THE DEVICE WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21161779&pc=LWP | 2124215-2025-01555 | 2024-12-22 05:00 | 2025-01-15 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXPLANTED AND REPLACED DUE TO ENTERING SAFETY MODE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER WAS RETURNED FOR ANALYSIS. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21139670&pc=LWP | 2124215-2025-01035 | 2024-12-16 05:00 | 2025-01-13 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | Dyspnea; Presyncope | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS PACEMAKER EXPERIENCED A PRE-SYNCOPAL EPISODE AND SHORTNESS OF BREATH. EVALUATION REVEALED THAT THE PACEMAKER WAS OPERATING IN SAFETY MODE. CONSEQUENTLY, THE PATIENT UNDERWENT A REVISION SURGERY DURING WHICH THE DEVICE WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21144677&pc=LWP | 2124215-2025-01017 | 2024-12-24 05:00 | 2025-01-13 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS SUSPECTED TO BE IN SAFETY MODE. THIS DEVICE WAS EXPLANTED AND SUCCESSFULLY REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21123782&pc=LWP | 2124215-2025-00639 | 2024-12-23 05:00 | 2025-01-10 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE DR | Fall; Syncope/Fainting | N970003 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS PACEMAKER EXPERIENCED A SYNCOPAL EPISODE. UPON INTERROGATION IT WAS DETERMINED THAT THIS PACEMAKER HAD ENTERED SAFETY MODE. THE PHYSICIAN OPTED TO EXPLANT AND REPLACE THIS DEVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE HAS NOT BEEN RETURNED.Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS PACEMAKER EXPERIENCED A SYNCOPAL EPISODE. UPON INTERROGATION IT WAS DETERMINED THAT THIS PACEMAKER HAD ENTERED SAFETY MODE. THE PHYSICIAN OPTED TO EXPLANT AND REPLACE THIS DEVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE HAS NOT BEEN RETURNED. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT DURING THE SYNCOPAL EPISODE THE PATIENT HAD A FALL WITHOUT SUSTAINED INJURY. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE HAS NOT BEEN RETURNED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21123943&pc=LWP | 2124215-2025-00734 | 2024-12-17 05:00 | 2025-01-10 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI EL DR | Dyspnea | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER DEVICE WAS FOUND IN SAFETY MODE. DURING THE FOLLOW UP VISIT, THE PATIENT REPORTED SYMPTOMS OF SLIGHT DYSPNEA ON EXERTION ALONG WITH ABNORMAL HEARTBEATS. THE ECG SHOWED THAT THE DEVICE WAS IN SAFETY MODE WITH NARROW QRS, DUE TO WHICH THE DEVICE WAS NOT ABLE TO BE INTERROGATED. SUBSEQUENTLY THIS DEVICE WAS EXPLANTED AND SUCCESSFULLY REPLACED. NO ADDITIONAL PATIENT ADVERSE EFFECTS WERE REPORTED. THIS DEVICE IS EXPECTED TO RETURN FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21125221&pc=LWP | 2124215-2025-00777 | 2024-12-12 05:00 | 2025-01-10 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER FELT INTO SAFETY MODE. TECHNICAL SERVICES (TS) WAS CONSULTED AND RECOMMENDED CHANGED OUT. THE DEVICE WAS EXPLANTED AND REPLACED. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21114365&pc=LWP | 2124215-2024-83435 | 2024-12-20 05:00 | 2025-01-09 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXPLANTED AS IT REVERTED TO SAFETY MODE OPERATION. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THE DEVICE IS EXPECTED TO BE RETURNED FOR ANALYSIS.Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXPLANTED AS IT REVERTED TO SAFETY MODE OPERATION. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THE DEVICE IS EXPECTED TO BE RETURNED FOR ANALYSIS. Manufacturer Narrative: AT THIS TIME, THE PRODUCT HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED, ANALYSIS WILL BE PERFORMED, AND THIS REPORT WILL BE UPDATED AT THAT TIME. Manufacturer Narrative: AT THIS TIME, THE PRODUCT HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED, ANALYSIS WILL BE PERFORMED, AND THIS REPORT WILL BE UPDATED AT THAT TIME.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21114762&pc=LWP | 2124215-2025-00505 | 2024-12-22 05:00 | 2025-01-09 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED AN ERROR MESSAGE DURING INTERROGATION. THE DEVICE WAS PART OF ADVISORY. LATER ON, THE DEVICE WAS EXPLANTED AND REPLACED AND ITS EXPECTED TO RETURN FOR ANALYSIS. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21119512&pc=LWP | 2124215-2024-83528 | 2024-12-18 05:00 | 2025-01-09 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER HAD ENTERED SAFETY MODE. THIS PACEMAKER REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THIS PACEMAKER HAD ENTERED SAFETY MODE. THIS PACEMAKER REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THIS PACEMAKER HAD ENTERED SAFETY MODE. THIS PACEMAKER REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION WAS RECEIVED AND THIS PACEMAKER HAS BEEN EXPLANTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PRODUCT HAS RETURNED FOR ANALYSIS. Manufacturer Narrative: THE DATE FOR THE DEVICE IMPLANT HAS BEEN REQUESTED FROM THE FIELD. IF AN UPDATE IS OBTAINED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE DATE FOR THE DEVICE IMPLANT HAS BEEN REQUESTED FROM THE FIELD. IF AN UPDATE IS OBTAINED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: THE DATE FOR THE DEVICE IMPLANT HAS BEEN REQUESTED FROM THE FIELD. IF AN UPDATE IS OBTAINED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21101341&pc=LWP | 2124215-2024-83029 | 2024-12-11 05:00 | 2025-01-08 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT DURING ROUTINE FOLLOW-UP, THIS PACEMAKER WAS FOUND TO BE IN SAFETY MODE. THE PATIENT WAS ASYMPTOMATIC AND NOT PACING DEPENDENT; HOWEVER, THEY UNDERWENT URGENT REPLACEMENT OF THE PACEMAKER ON THE (B)(6) 2024, WITHOUT COMPLICATION. THE PACEMAKER IS EXPECTED TO BE RETURNED FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21106873&pc=LWP | 2124215-2025-00319 | 2024-12-12 05:00 | 2025-01-08 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER DISPLAYED SAFETY MODE UPON INTERROGATION IN CLINIC. THIS PACEMAKER REMAINS IN SERVICE AT THIS TIME. NO ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THIS PACEMAKER DISPLAYED SAFETY MODE UPON INTERROGATION IN CLINIC. THIS PACEMAKER REMAINS IN SERVICE AT THIS TIME. NO ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION WAS RECEIVED THAT THIS PACEMAKER WAS EXPLANTED AND REPLACED WITH A NEW PACEMAKER. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21108167&pc=LWP | 2124215-2025-00390 | 2024-12-13 05:00 | 2025-01-08 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE DR | Device Overstimulation of Tissue | N970003 | 1 | Event Description: IT WAS REPORTED THAT DURING A CLINIC VISIT THIS PACEMAKER WAS INTERROGATED AND FOUND TO BE IN SAFETY MODE. THE PATIENT ALSO HAD REPORTED EXPERIENCING UNWANTED POCKET STIMULATION. A DEVICE REPLACEMENT PROCEDURE WAS SCHEDULED AND LATER PERFORMED. THIS PACEMAKER WAS EXPLANTED AND REPLACED WITH A NEW DEVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THE DEVICE WAS RETAINED BY THE HOSPITAL.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21054937&pc=LWP | 2124215-2024-82807 | 2024-12-13 05:00 | 2025-01-02 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS FOUND TO BE IN SAFETY MODE. THIS DEVICE WAS EXPLANTED AND REPLACED WITH A NEW PACEMAKER. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21056021&pc=LWP | 2124215-2024-70534 | 2024-10-16 04:00 | 2025-01-02 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT DURING A ROUTINE FOLLOW UP, THIS PACEMAKER WAS FOUND TO BE IN SAFETY MODE. A REPLACEMENT PROCEDURE WAS PERFORMED, THIS DEVICE WAS EXPLANTED AND REPLACED WITH A NEW DEVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THE DEVICE IS EXPECTED TO RETURN FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21061294&pc=LWP | 2124215-2025-00080 | 2024-12-20 05:00 | 2025-01-02 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE PACEMAKER WAS EXPLANTED AND REPLACED DUE TO SAFETY MODE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21050406&pc=LWP | 2124215-2024-82705 | 2024-12-19 05:00 | 2024-12-31 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE DR | No Clinical Signs, Symptoms or Conditions | N970003/S167 | 1 | Event Description: IT WAS REPORTED THAT DEVICE INTERROGATION REVEALED THIS PACEMAKER ENTERED SAFETY MODE, AND DEVICE REPLACEMENT WAS SCHEDULED. NO ADVERSE PATIENT EFFECTS OR INTERVENTIONS WERE REPORTED AT THIS TIME. IT WAS NOTED THAT THE PATIENT WAS PACER DEPENDENT.Event Description: IT WAS REPORTED THAT DEVICE INTERROGATION REVEALED THIS PACEMAKER ENTERED SAFETY MODE, AND DEVICE REPLACEMENT WAS SCHEDULED. NO ADVERSE PATIENT EFFECTS OR INTERVENTIONS WERE REPORTED AT THIS TIME. IT WAS NOTED THAT THE PATIENT WAS PACER DEPENDENT. ADDITIONAL INFORMATION RECEIVED REPORTED THAT THE PACEMAKER WAS EXPLANTED AND REPLACED DUE TO ENTERING SAFETY MODE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. PRODUCT RETURN WAS REQUESTED. Manufacturer Narrative: THE LOCAL AREA FIELD REPRESENTATIVE WAS CONTACTED FOR ADDITIONAL INFORMATION REGARDING EVENT CAUSE/RESOLUTION AND PRODUCT RETURN STATUS. DEVICE REPLACEMENT WAS CONFIRMED, AND PRODUCT RETURN IS IN PROGRESS. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. Manufacturer Narrative: THE LOCAL AREA FIELD REPRESENTATIVE WAS CONTACTED FOR ADDITIONAL INFORMATION REGARDING EVENT CAUSE/RESOLUTION AND PRODUCT RETURN STATUS. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21027916&pc=LWP | 2124215-2024-82156 | 2024-12-05 05:00 | 2024-12-27 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER FELT INTO SAFETY MODE. THE DEVICE WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21010003&pc=LWP | 2124215-2024-81680 | 2024-12-16 05:00 | 2024-12-24 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER HAD ENTERED SAFETY MODE. THE PHYSICIAN OPTED TO EXPLANT AND REPLACE THIS DEVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE HAS NOT BEEN RETURNED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21001544&pc=LWP | 2124215-2024-81466 | 2024-12-16 05:00 | 2024-12-23 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE DR | Discomfort; Unspecified Heart Problem | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER FELT INTO SAFETY MODE. TECHNICAL SERVICES (TS) WAS CONSULTED AND RECOMMENDED CHANGED OUT. THE PATIENT EXPERIENCED DISCOMFORT AND PALPITATIONS. THE DEVICE WAS EXPLANTED AND REPLACED. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21002014&pc=LWP | 2124215-2024-81496 | 2024-12-09 05:00 | 2024-12-23 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT DURING A CLINIC VISIT, THIS PACEMAKER WAS FOUND TO BE IN SAFETY MODE. A DEVICE REPLACEMENT PROCEDURE WAS LATER PERFORMED. THE DEVICE WAS SUCCESSFULLY EXPLANTED AND REPLACED WITH A NEW DEVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20985156&pc=LWP | 2124215-2024-80931 | 2024-12-05 05:00 | 2024-12-20 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | Discomfort | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER HAD ENTERED INTO SAFETY MODE AND THAT THE PATIENT NOTICED A CHANGE IN THE PACING OUTPUT. THE DEVICE WAS SUBSEQUENTLY EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE HAS BEEN RETURNED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20973751&pc=LWP | 2124215-2024-80477 | 2024-11-25 05:00 | 2024-12-19 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT DURING A RETROSPECTIVE REVIEW OF DEVICE DATA IT WAS IDENTIFIED THAT THIS PACEMAKER WAS FOUND TO BE IN SAFETY MODE. THIS DEVICE WAS THEN SUCCESFULLY EXPLANTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT DURING A RETROSPECTIVE REVIEW OF DEVICE DATA IT WAS IDENTIFIED THAT THIS PACEMAKER WAS FOUND TO BE IN SAFETY MODE. NO ADDITIONAL DETAILS WERE PROVIDED. AT THIS POINT, THE DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20980936&pc=LWP | 2124215-2024-80768 | 2024-12-16 05:00 | 2024-12-19 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXPLANTED AND REPLACED DUE TO ENTERING SAFETY MODE. OTHER THAN THE SURGICAL PROCEDURE, NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER HAS NOT BEEN RETURNED FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20972696&pc=LWP | 2124215-2024-80511 | 2024-12-11 05:00 | 2024-12-18 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS FOUND TO BE IN SAFETY MODE. TECHNICAL SERVICES (TS) WAS CONSULTED AND RECOMMENDED DEVICE REPLACEMENT. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS FOUND TO BE IN SAFETY MODE. TECHNICAL SERVICES (TS) WAS CONSULTED AND RECOMMENDED DEVICE REPLACEMENT. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT THIS PATIENT UNDERWENT A REVISION SURGERY IN WHICH THIS PACEMAKER WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER HAS NOT BEEN RECEIVED FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20940308&pc=LWP | 2124215-2024-79384 | 2024-12-02 05:00 | 2024-12-16 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE EL DR | No Clinical Signs, Symptoms or Conditions | N970003/S167 | 1 | Event Description: IT WAS REPORTED THAT DURING ROUTINE FOLLOW UP, THIS PACEMAKER DEVICE WAS FOUND TO BE IN SAFETY MODE. A DEVICE REPLACEMENT PROCEDURE WAS PERFORMED. THE DEVICE WAS SUCCESSFULLY REMOVED AND REPLACED WITH A NEW DEVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20947389&pc=LWP | 2124215-2024-79565 | 2024-12-05 05:00 | 2024-12-16 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE EL DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS FOUND TO BE IN SAFETY MODE DURING A ROUTINE CHECK. IT WAS RECOMMENDED THAT THIS DEVICE BE REPLACED. THE DEVICE WAS SUBSEQUENTLY EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED OUTSIDE OF THIS REVISION PROCEDURE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20915125&pc=LWP | 2124215-2024-78711 | 2024-11-28 05:00 | 2024-12-12 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE DR | Pain; Discomfort | N970003 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH THIS PACEMAKER SYSTEM PRESENTED TO THE EMERGENCY ROOM (ER) WITH SYMPTOMS OF PAIN AND DISCOMFORT. THE PACEMAKER HAD ENTERED SAFETY MODE AND INTERROGATION REVEALED THE FOLLOWING CODES: 0X0 0X0 0X0. IT WAS DETERMINED THAT THE PATIENTS SYMPTOMS WERE SECONDARY TO THE SAFETY MODE SWITCH TO UNIPOLAR PACING. THE PATIENT WAS TRANSFERRED TO A DIFFERENT MEDICAL CENTER VIA AMBULANCE, AND THE DEVICE WAS SUCCESSFULLY EXPLANTED AND REPLACED. LEAD NUMBERS WERE WITHIN NORMAL LIMITS, AND THE PATIENT WAS DISCHARGED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER WAS RETURNED FOR ANALYSIS. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20909912&pc=LWP | 2124215-2024-78550 | 2024-11-28 05:00 | 2024-12-11 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER HAD ENTERED SAFETY MODE. THE PHYSICIAN OPTED TO EXPLANT AND REPLACE THIS DEVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE HAS NOT BEEN RETURNED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21279791&pc=LWP | 2124215-2025-04926 | 2025-01-21 05:00 | 2025-01-30 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings; Battery Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS DEVICE HAD ENTERED SAFETY MODE DESPITE HAVING A REMAINING LONGEVITY OF 1.5 YEARS 3 MONTHS AGO. THE DEVICE WAS EXPLANTED AND REPLACED WITHOUT FURTHER INCIDENT. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: AS NO FURTHER INFORMATION CONCERNING THIS REPORT IS EXPECTED, OUR INVESTIGATION IS COMPLETE. THIS INVESTIGATION WILL BE UPDATED SHOULD FURTHER INFORMATION BE PROVIDED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20984573&pc=LWP | 2124215-2024-80889 | 2024-11-21 05:00 | 2024-12-19 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings; Defective Device; Interrogation Problem | ACCOLADE EL DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXPLANTED DUE TO PRODUCT PERFORMANCE ISSUE AFTER NINE YEARS OF BEING IMPLANTED. IT WAS REPLACED AND RETAINED BY CUSTOMER. NO ADDITIONAL PATIENT ADVERTS EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXPLANTED DUE TO PRODUCT PERFORMANCE ISSUE AFTER NINE YEARS OF BEING IMPLANTED. IT WAS REPLACED AND RETAINED BY CUSTOMER. NO ADDITIONAL PATIENT ADVERTS EFFECTS WERE REPORTED. SUBSEQUENT ADDITIONAL INFORMATION WAS PROVIDED REPORTED THAT DURING A ROUTINE FOLLOW UP, THE PACEMAKER WAS NOTED TO HAVE DISPLAYED A RED ALERT. DEVICE INTERROGATION WAS ATTEMPTED, HOWEVER WAS UNSUCCESSFUL. THE PHYSICIAN PERFORMED A REPLACEMENT PROCEDURE, THE PACEMAKER WAS THEN EXPLANTED AND REPLACED WITH A NEW DEVICE SUCCESSFULLY. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THE DEVICE WAS RETURNED TO THE FACILITY FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21161821&pc=LWP | 2124215-2025-01596 | 2024-12-31 05:00 | 2025-01-15 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings; Impedance Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXPLANTED DUE TO IMPEDANCE ISSUES. A NEW DEVICE WAS SUCCESSFULLY IMPLANTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE HAS BEEN RECEIVED FOR ANALYSIS.Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXPLANTED DUE TO IMPEDANCE ISSUES. A NEW DEVICE WAS SUCCESSFULLY IMPLANTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE HAS BEEN RECEIVED FOR ANALYSIS. FURTHER INFORMATION WAS RECEIVED THAT THE DEVICE WAS FOUND TO BE IN SAFETY MODE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20994549&pc=LWP | 2124215-2024-81238 | 2024-12-03 05:00 | 2024-12-20 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings; Inappropriate or Unexpected Reset | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER ENTERED SAFETY MODE AND RECORDED CODES 0XA 0XA 0XA, INDICATIVE OF AN OSCILLATOR MISMATCH AND DEVICE RESET. SUBSEQUENTLY, THIS PACEMAKER WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED, AND IT WAS NOTED THAT THE PATIENT WAS NOT PACER DEPENDENT. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. PRODUCT RETURN WAS REQUESTED. Manufacturer Narrative: THE LOCAL AREA FIELD REPRESENTATIVE WAS CONTACTED FOR ADDITIONAL INFORMATION REGARDING PRODUCT RETURN STATUS. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20958600&pc=LWP | 2124215-2024-79986 | 2024-12-05 05:00 | 2024-12-17 05:00 | Injury | Incorrect, Inadequate or Imprecise Result or Readings; Telemetry Discrepancy | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS UNABLE TO BE INTERROGATED WITH RADIO FREQUENCY USING THE COMMUNICATOR, THE PATIENT WAS BROUGHT INTO CLINIC, AND THE DEVICE WAS FOUND TO HAVE REVERTED TO SAFETY MODE. THIS DEVICE WAS EXPLANTED AND SUCCESSFULLY REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21251955&pc=LPH | 0002249697-2025-00118 | 2024-12-31 05:00 | 2025-01-28 05:00 | Injury | Material Erosion | ACCOLADE 132 SIZE 4.5 | Ambulation Difficulties | K233498 | 1 | Event Description: RIGHT TOTAL HIP REVISION DUE TO EXCESSIVE FEMORAL STEM NECK TRUNNION WEAR. REVISED STEM, HEAD, ACETABULAR POLY LINER. Manufacturer Narrative: REPORTED EVENT: AN EVENT REGARDING WEAR INVOLVING AN ACCOLADE STEM WAS REPORTED. THE EVENT WAS CONFIRMED VIA MEDIAL REVIEW AND PROVIDED PHOTOGRAPHS. METHOD & RESULTS: PRODUCT EVALUATION AND RESULTS: THE DEVICE WAS NOT RETURNED FOR REVIEW; HOWEVER, PHOTOGRAPHS WERE PROVIDED. THE PHOTOGRAPH CONFIRMS THE STEM TRUNNION WITH SIGNIFICANT MATERIAL LOSS OF THE PROXIMAL AND SUPERIOR ASPECTS OF THE TRUNNION. CLINICIAN REVIEW: A REVIEW OF THE PROVIDED MEDICAL RECORDS BY A CLINICAL CONSULTANT INDICATED: "THE INTRA-OPERATIVE PHOTOGRAPH CONFIRMS REMODELING OF THE FEMORAL STEM TRUNNION WITH SIGNIFICANT MATERIAL LOSS OF THE PROXIMAL AND SUPERIOR ASPECTS OF THE TRUNNION. THE ROOT CAUSE OF THIS MECHANICAL COMPLICATION CANNOT BE DETERMINED FROM THE DOCUMENTATION PROVIDED. SHOULD FURTHER DOCUMENTATION BECOME AVAILABLE I WOULD BE HAPPY TO FURTHER THIS INVESTIGATION." PRODUCT HISTORY REVIEW: REVIEW OF THE DEVICE HISTORY RECORDS INDICATE DEVICES WERE MANUFACTURED AND ACCEPTED INTO FINAL STOCK WITH NO RELEVANT REPORTED DISCREPANCIES. COMPLAINT HISTORY REVIEW: THERE HAVE BEEN NO OTHER SIMILAR EVENTS FOR LOT REFERENCED. CONCLUSIONS: IT WAS REPORTED THAT THE PATIENT WAS REVISED DUE TO STEM TRUNNION WEAR. A REVIEW OF THE PROVIDED MEDICAL RECORDS BY A CLINICAL CONSULTANT INDICATED: "THE INTRA-OPERATIVE PHOTOGRAPH CONFIRMS REMODELING OF THE FEMORAL STEM TRUNNION WITH SIGNIFICANT MATERIAL LOSS OF THE PROXIMAL AND SUPERIOR ASPECTS OF THE TRUNNION. THE ROOT CAUSE OF THIS MECHANICAL COMPLICATION CANNOT BE DETERMINED FROM THE DOCUMENTATION PROVIDED. SHOULD FURTHER DOCUMENTATION BECOME AVAILABLE I WOULD BE HAPPY TO FURTHER THIS INVESTIGATION." NO FURTHER INVESTIGATION FOR THIS EVENT IS POSSIBLE AT THIS TIME. IF DEVICES AND / OR ADDITIONAL INFORMATION BECOME AVAILABLE TO INDICATE FURTHER EVALUATION IS WARRANTED, THIS RECORD WILL BE REOPENED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20978447&pc=LPH | 0002249697-2024-01750 | 2022-08-08 04:00 | 2024-12-19 05:00 | Injury | Material Erosion | ACCOLADE (127 DEG) SIZE 3.5 ACCOLADE (127 DEG) SIZE 3.5 | Metal Related Pathology | K121308 | 1 | Event Description: IT WAS REPORTED BY COUNSEL THAT CLAIMANT UNDERWENT RIGHT THA ON (B)(6) 2011 AND WAS IMPLANTED WITH AN LFIT ANATOMIC V40 FEMORAL HEAD AND ACCOLADE TMZF HIP STEM. HIS RIGHT HIP WAS REVISED ON (B)(6) 2022 ALLEGEDLY DUE TO ELEVATED METAL LEVELS IN HIS BLOOD AND SUSPECTED ALVAL. Manufacturer Narrative: IT WAS NOTED THAT THE DEVICE IS NOT AVAILABLE FOR EVALUATION. IF ADDITIONAL INFORMATION IS RECEIVED, IT WILL BE PROVIDED IN A SUPPLEMENTAL REPORT UPON COMPLETION OF THE INVESTIGATION. REVIEW OF THE DEVICE HISTORY RECORDS INDICATE DEVICES WERE MANUFACTURED AND ACCEPTED INTO FINAL STOCK WITH NO REPORTED DISCREPANCIES. THERE HAVE BEEN NO OTHER EVENTS FOR THE LOT REFERENCED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21261961&pc=LWP | 2124215-2025-04345 | 2025-01-03 05:00 | 2025-01-29 05:00 | Injury | Migration | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER WAS REPOSITIONED DUE TO DEVICE MOVING FROM SUBCUTANEOUS TO SUBMUSCULAR. THIS DEVICE REMAINS IN SERVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21110615&pc=LWP | 2124215-2025-00373 | 2024-12-05 05:00 | 2025-01-09 05:00 | Injury | Migration | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE IMPLANTABLE PACEMAKER WAS REPOSITIONED DUE TO DEVICE MIGRATED TO THE ARMPIT. THIS DEVICE REMAINS IN SERVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21051155&pc=LWP | 2124215-2024-82730 | 2024-12-05 05:00 | 2025-01-01 05:00 | Injury | Migration | ACCOLADE MRI DR | Undesired Nerve Stimulation | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER DEVICE HAS MIGRATED AFTER IMPLANT WHICH DESPITE BEING SUTURED TO THE PECTORALIS DUE TO SUSPECTED TO BE A RESULT OF THE PATIENT'S BREAST SIZE. ALSO, PATIENT EXPERIENCED PHRENIC NERVE STIMULATION. THE DEVICE THERAPY WAS TURNED OR PROGRAMMED OFF AND DEVICE WILL BE REPLACED WITH A LEADLESS PACEMAKER. AS OF THIS TIME, THIS DEVICE REMAINS IN SERVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21009080&pc=LWP | 2124215-2024-81521 | 2024-11-24 05:00 | 2024-12-23 05:00 | Injury | Over-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM WAS FOUND TO EXHIBIT OVERSENSING ON THE RIGHT VENTRICULAR (RV) LEAD. A REVISION PROCEDURE WAS PERFORMED, THE PACEMAKER AND RV LEAD WERE EXPLANTED AND REPLACED WITH A NEW DEVICE AND LEAD. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM WAS FOUND TO EXHIBIT OVERSENSING ON THE RIGHT VENTRICULAR (RV) LEAD. A REVISION PROCEDURE WAS PERFORMED, THE PACEMAKER AND RV LEAD WERE EXPLANTED AND REPLACED WITH A NEW DEVICE AND LEAD. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION WAS RECEIVED FROM THE FIELD REPRESENTATIVE THAT REPORTED THAT THE OVERSENSING ISSUES WAS OBSERVED ON THE RIGHT ATRIAL (RA) LEAD AND NOT THE RV LEAD AND THE LEAD REVISION WAS PERFORMED ON THE RA LEAD DUE TO LEAD INTEGRITY ISSUES OBSERVED ON THE RA LEAD. THE PACEMAKER AND RA LEAD WERE SUCCESSFULLY EXPLANTED AND REPLACED WITH NEW DEVICE AND LEAD. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: CORRECTION: B5:-DESCRIBE EVENT OR PROBLEM.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21211597&pc=LWP | 2124215-2025-02985 | 2025-01-07 05:00 | 2025-01-22 05:00 | Injury | Over-Sensing; Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI EL DR | Device Overstimulation of Tissue | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS FOUND TO BE IN SAFETY MODE. THE PATIENT EXPERIENCED MUSCLE STIMULATION AND INTERMITTENT OVERSENSING WAS ALSO OBSERVED. A DEVICE CHANGEOUT PROCEDURE WAS PERFORMED AND IT WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE HAS BEEN RECEIVED FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21266148&pc=LWP | 2124215-2025-04478 | 2025-01-14 05:00 | 2025-01-29 05:00 | Injury | Over-Sensing; Pacing Problem; Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI EL DR | Cardiac Arrest; Dizziness; Asystole | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS PACEMAKER EXPERIENCED LIGHT-HEADEDNESS AND VISITED THE EMERGENCY ROOM. SUBSEQUENTLY, THE PACEMAKER WAS FOUND TO BE OPERATING IN SAFETY MODE. UPON REVIEWING THE EPISODE, IT WAS NOTED THAT THE PATIENT HAD ALSO EXPERIENCED A CARDIAC ARREST LASTING APPROXIMATELY SEVEN SECONDS DUE TO OVERSENSING AND PACING INHIBITION. CONSEQUENTLY, THE PATIENT WAS HOSPITALIZED, AND A TEMPORARY DEVICE WAS INSERTED. THE PATIENT LATER UNDERWENT REVISION SURGERY, DURING WHICH THE PACEMAKER WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER HAS NOT BEEN RECEIVED FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21223754&pc=LWP | 2124215-2025-03228 | 2025-01-09 05:00 | 2025-01-23 05:00 | Injury | Over-Sensing; Pacing Problem; Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI EL DR | Dizziness; Presyncope | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM ENTERED SAFETY MODE. THE PACEMAKER EXHIBITED INTERMITTENT PAUSES IN PACING, AND THE PATIENT WAS EXPERIENCING LIGHTHEADEDNESS AND DIZZINESS. THE PATIENT WAS ADMITTED TO THE HOSPITAL AND EXTERNAL PACING PATCHES UNTIL THE DEVICE REPLACEMENT COULD BE PERFORMED. ADDITIONALLY, THE PATIENT'S LEFT ARM WAS IMMOBILIZED IN A SLING TO MINIMIZE MOTION AND MYOPOTENTIAL OVERSENSING. THE NEXT DAY, THE PACEMAKER WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER WAS RETURNED FOR ANALYSIS. Manufacturer Narrative: THE PRODUCT HAS BEEN RECEIVED FOR ANALYSIS. THIS REPORT WILL BE UPDATED UPON COMPLETION OF ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21010004&pc=LWP | 2124215-2024-81677 | 2024-12-05 05:00 | 2024-12-24 05:00 | Injury | Over-Sensing; Pacing Problem; Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | Dizziness | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER HAD ENTERED SAFETY MODE, WHICH CAUSED OVERSENSING AND PACING INHIBITION LEADING TO LIGHTHEADEDNESS SYMPTOMS ON THE PATIENT. THE PHYSICIAN OPTED TO EXPLANT AND REPLACE THIS DEVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE HAS NOT BEEN RETURNED.Event Description: IT WAS REPORTED THAT THIS PACEMAKER HAD ENTERED SAFETY MODE, WHICH CAUSED OVERSENSING AND PACING INHIBITION LEADING TO LIGHTHEADEDNESS SYMPTOMS ON THE PATIENT. THE PHYSICIAN OPTED TO EXPLANT AND REPLACE THIS DEVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE HAS BEEN RETURNED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21015235&pc=LWP | 2124215-2024-81861 | 2024-04-03 04:00 | 2024-12-25 05:00 | Injury | Pacing Problem; High Capture Threshold | ACCOLADE MRI DR | Unspecified Infection | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER SHOWED A PULSE RATE OF 50 THE SET LOW RATE WAS AT 60. ADDITIONAL INFORMATION RECEIVED INDICATES THAT PACING FAILURE WAS SUSPECTED DUE TO INCREASED PACING THRESHOLD. THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION RECEIVED INDICATES THAT THE DEVICE WAS EXPLANTED AND REPLACED. THIS DEVICE WAS DISCARDED IN THE FACILITY DUE TO INFECTION. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20967878&pc=LWP | 2124215-2024-80278 | 2024-12-10 05:00 | 2024-12-18 05:00 | Injury | Pacing Problem; Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE DR | Pain; Discomfort | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER ENTERED SAFETY MODE, WHICH CAUSED DISCOMFORT AND PAIN TO THE PATIENT SINCE THE DEVICE SWITCHED TO UNIPOLAR PACING AND THERE WAS A LACK OF THE REQUIRED ATRIAL PACING UNDER THE SAFETY MODE SETTINGS. SUBSEQUENTLY, THIS DEVICE WAS EXPLANTED. A NEW DEVICE WAS IMPLANTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THIS PACEMAKER ENTERED SAFETY MODE, WHICH CAUSED DISCOMFORT AND PAIN TO THE PATIENT SINCE THE DEVICE SWITCHED TO UNIPOLAR PACING AND THERE WAS A LACK OF THE REQUIRED ATRIAL PACING UNDER THE SAFETY MODE SETTINGS. SUBSEQUENTLY, THIS DEVICE WAS EXPLANTED. A NEW DEVICE WAS IMPLANTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: CORRECTED FIELDS: H6 IMPACT CODES.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20967890&pc=LWP | 2124215-2024-80297 | 2024-12-11 05:00 | 2024-12-18 05:00 | Injury | Pacing Problem; Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS FOUND TO BE IN SAFETY MODE. A DEVICE CHANGEOUT PROCEDURE WAS PERFORMED AND IT WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21207929&pc=LWP | 2124215-2025-02849 | 2025-01-03 05:00 | 2025-01-22 05:00 | Injury | Pocket Stimulation; Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE EL DR | Dyspnea; Discomfort | N970003 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS PACEMAKER FELT A PULSING FROM THE DEVICE ALONG WITH SHORTNESS OF BREATH (SOB) AND DISCOMFORT. THE PATIENT PRESENTED TO THE EMERGENCY ROOM (ER) WHERE IT WAS FOUND THAT THIS DEVICE WAS OPERATING IN SAFETY MODE. SUBSEQUENTLY, THIS PACEMAKER WAS EXPLANTED AND REPLACED WITH A NEW ONE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21242100&pc=LWP | 2124215-2025-03781 | 2024-12-31 05:00 | 2025-01-27 05:00 | Injury | Premature Discharge of Battery | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER DEVICE WAS SURGICALLY EXPLANTED DUE TO ADVISORY OR RECALL. THIS DEVICE WAS REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21212287&pc=LWP | 2124215-2025-03000 | 2024-12-26 05:00 | 2025-01-22 05:00 | Injury | Premature Discharge of Battery | ACCOLADE EL DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER DEVICE EXHIBITED PREMATURE BATTERY DEPLETION (PBD). THE IMPLANTABLE PACEMAKER DEVICE WAS EXPLANTED AND A NEW DEVICE WITH A DIFFERENT MODEL WAS PLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21214413&pc=LWP | 2124215-2025-03088 | 2024-12-27 05:00 | 2025-01-22 05:00 | Injury | Premature Discharge of Battery | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER EXHIBITED PREMATURE BATTERY DEPLETION (PBD). THIS DEVICE WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21197454&pc=LWP | 2124215-2025-02577 | 2024-12-24 05:00 | 2025-01-21 05:00 | Injury | Premature Discharge of Battery | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS IMPLANTABLE PACEMAKER DEVICE EXHIBITED PREMATURE BATTERY DEPLETION. THIS DEVICE WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21199760&pc=LWP | 2124215-2025-02667 | 2024-12-26 05:00 | 2025-01-21 05:00 | Injury | Premature Discharge of Battery | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER DEVICE EXHIBITED PREMATURE BATTERY DEPLETION. THIS DEVICE WAS EXPLANTED AND REPLACED DUE TO ADVISORY OR RECALL. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21204495&pc=LWP | 2124215-2025-02805 | 2024-12-22 05:00 | 2025-01-21 05:00 | Injury | Premature Discharge of Battery | ACCOLADE SR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER WAS EXPLANTED DUE TO PREMATURE BATTERY DEPLETION. A NEW DEVICE WAS IMPLANTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21158123&pc=LWP | 2124215-2025-01450 | 2024-12-16 05:00 | 2025-01-15 05:00 | Injury | Premature Discharge of Battery | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER DEVICE WAS IMPLANTED FOR THREE YEARS AND WAS EXPLANTED. THIS DEVICE WAS REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21164055&pc=LWP | 2124215-2025-01653 | 2024-12-17 05:00 | 2025-01-15 05:00 | Injury | Premature Discharge of Battery | ACCOLADE MRI DR | Bradycardia | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER RECORDED A CODE 1003 INDICATIVE OF BATTERY VOLTAGE TOO LOW FOR THE PROJECTED REMAINING CAPACITY. ALSO, PATIENT HAD HEART RATE IN THE 30'S. ALSO, BATTERY DEPLETION WAS LIKELY RESULT OF OUTPUTS HAD BEEN LEFT HIGHER FOR MOST OF DEVICE LIFE. THIS DEVICE WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21152267&pc=LWP | 2124215-2025-01335 | 2024-12-18 05:00 | 2025-01-14 05:00 | Injury | Premature Discharge of Battery | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER DEVICE EXHIBITED PREMATURE BATTERY DEPLETION. THIS DEVICE WAS EXPLANTED AND REPLACED WITH THE SAME MODEL. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21123734&pc=LWP | 2124215-2025-00695 | 2023-09-21 04:00 | 2025-01-10 05:00 | Injury | Premature Discharge of Battery | ACCOLADE DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER EXHIBITED PREMATURE BATTERY DEPLETION (PBD). THIS DEVICE WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21114442&pc=LWP | 2124215-2025-00467 | 2024-12-11 05:00 | 2025-01-09 05:00 | Injury | Premature Discharge of Battery | ACCOLADE MRI DR | Bradycardia | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER WAS SURGICALLY EXPLANTED DUE TO PREMATURE BATTERY DEPLETION. IN ADDITION, THE PATIENT WAS IN THE HOSPITAL AND TELEMETRY SHOWED THAT THE HEART RATE WAS IN FORTIES WITH NO PACE SPIKES. THIS DEVICE WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21105131&pc=LWP | 2124215-2025-00267 | 2023-01-10 05:00 | 2025-01-08 05:00 | Injury | Premature Discharge of Battery | ACCOLADE EL DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS PACEMAKER WAS SUSPECTED TO BE DEPLETING PREMATURELY. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS CONFIRMED THE BATTERY APPEARED TO BE DEPLETING NORMALLY. DEVICE REPLACEMENT WAS RECOMMENDED. THE DEVICE REMAINS IN SERVICE AND THERE IS NO EVIDENCE TO SUGGEST A DEVICE REPLACEMENT PROCEDURE HAS BEEN SCHEDULED AT THIS TIME. NO ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION WAS RECEIVED, THERE WERE CONCERNS OF PREMATURE BATTERY DEPLETION (PBD) FOR THIS PACEMAKER. THIS PACEMAKER WAS EXPLANTED AND SUCCESSFULLY REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21056054&pc=LWP | 2124215-2024-70572 | 2024-08-23 04:00 | 2025-01-02 05:00 | Injury | Premature Discharge of Battery | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER DEVICE RECORDED A CODE 1003 INDICATIVE OF BATTERY VOLTAGE TOO LOW FOR THE PROJECTED REMAINING CAPACITY. BOSTON SCIENTIFIC TECHNICAL SERVICE (TS) REVIEWED AND DISCUSSED THAT THE DEVICE CURRENT DRAIN WAS GETTING WORSE AS TIME PROGRESSED. FURTHERMORE, THIS DEVICE WAS EXPLANTED FOR PREMATURE BATTERY DEPLETION (PBD). NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER DEVICE RECORDED A CODE 1003 INDICATIVE OF BATTERY VOLTAGE TOO LOW FOR THE PROJECTED REMAINING CAPACITY. BOSTON SCIENTIFIC TECHNICAL SERVICE (TS) REVIEWED AND DISCUSSED THAT THE DEVICE CURRENT DRAIN WAS GETTING WORSE AS TIME PROGRESSED. FURTHERMORE, THIS DEVICE WAS EXPLANTED FOR PREMATURE BATTERY DEPLETION (PBD). NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY, THIS DEVICE WAS THOROUGHLY INSPECTED AND ANALYZED. REVIEW OF DEVICE MEMORY CONFIRMED THAT A LOW VOLTAGE ALERTCODE 1003) WAS RECORDED. A HIGH CURRENT DRAIN WAS DETECTED, BUT THE BATTERY STILL SUPPORTED FULL DEVICE FUNCTION. DETAILED ANALYSIS CONFIRMED THE IDENTIFIED HIGH CURRENT DRAIN WAS A RESULT OF COMPROMISED CAPACITORS. THIS RESULTED IN THE OBSERVED PREMATURE BATTERY DEPLETION. IT WAS DETERMINED THAT THE CAPACITORS WERE COMPROMISED DUE TO THE PRESENCE OF EXCESS HYDROGEN IN THE DEVICE CASE. BOSTON SCIENTIFIC ISSUED A FIELD SAFETY NOTICE REGARDING A SUBSET OF DEVICES IN THE ACCOLADE PACEMAKER FAMILY THAT HAVE A POTENTIAL OF EXHIBITING THIS BEHAVIOR. THIS PARTICULAR DEVICE IS INCLUDED IN THE HYDROGEN INDUCED PREMATURE DEPLETION ADVISORY POPULATION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20989945&pc=LWP | 2124215-2024-81086 | 2024-12-13 05:00 | 2024-12-20 05:00 | Injury | Premature Discharge of Battery | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS IMPLANTABLE DEVICE WAS SUSPECTED TO BE DEPLETING PREMATURELY. ADDITIONAL INFORMATION PROVIDED CONFIRMED THAT PRODUCT PERFORMANCE ISSUE WAS THE REASONING FOR EXPLANT. THIS DEVICE WAS REPLACED AND RETURNED TO BOSTON SCIENTIFIC FOR ANALYSIS. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20976685&pc=LWP | 2124215-2024-80562 | 2024-12-09 05:00 | 2024-12-19 05:00 | Injury | Premature Discharge of Battery | ACCOLADE MRI SR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER HAD BEEN NEARING ELECTIVE REPLACEMENT INDICATOR (ERI) BATTERY STATUS. HOWEVER, AT THE CURRENT DEVICE FOLLOW UP, A CODE 1003 WAS OBSERVED. DUE TO THE ERROR MESSAGE, PREMATURE BATTERY DEPLETION WAS SUSPECTED AND THE DEVICE WAS EXPLANTED AND REPLACED THAT DAY. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THE EXPLANTED DEVICE WAS RETURNED FOR ANALYSIS. Manufacturer Narrative: THE PRODUCT HAS BEEN RECEIVED FOR ANALYSIS. THIS REPORT WILL BE UPDATED UPON COMPLETION OF ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20967863&pc=LWP | 2124215-2024-80223 | 2024-12-11 05:00 | 2024-12-18 05:00 | Injury | Premature Discharge of Battery | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER DEVICE EXHIBITED PREMATURE BATTERY DEPLETION (PBD). SUBSEQUENTLY, THIS DEVICE WAS EXPLANTED AND SUCCESSFULLY REPLACED. NO ADDITIONAL PATIENT ADVERSE EFFECTS WERE REPORTED. THIS DEVICE IS EXPECTED TO RETURN FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20921387&pc=LWP | 2124215-2024-78885 | 2024-11-06 05:00 | 2024-12-12 05:00 | Injury | Premature Discharge of Battery | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS IMPLANTABLE DEVICE WAS SUSPECTED TO BE DEPLETING PREMATURELY. THIS DEVICE WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: AS NO FURTHER INFORMATION CONCERNING THIS REPORT IS EXPECTED, OUR INVESTIGATION IS COMPLETE. THIS INVESTIGATION WILL BE UPDATED SHOULD FURTHER INFORMATION BE PROVIDED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21218625&pc=LWP | 2124215-2025-03132 | 2025-01-12 05:00 | 2025-01-23 05:00 | Injury | Premature Discharge of Battery; Battery Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS PACEMAKER WAS SUSPECTED TO BE DEPLETING PREMATURELY. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS CONFIRMED THE BATTERY APPEARED TO BE DEPLETING MORE QUICKLY THAN EXPECTED. DEVICE REPLACEMENT WAS RECOMMENDED. THIS DEVICE WAS EXPLANTED AND REPLACED. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21204240&pc=LWP | 2124215-2025-02788 | 2022-12-09 05:00 | 2025-01-21 05:00 | Injury | Premature Discharge of Battery; Battery Problem | ACCOLADE EL DR | No Clinical Signs, Symptoms or Conditions | N970003/S167 | 1 | Event Description: IT WAS REPORTED THAT IMPLANTABLE PACEMAKER HAD ONLY TWO YEARS REMAINING. DATA ANALYSIS NOTED NO SIGNS OF PREMATURE BATTERY DEPLETION (PBD) AND NOR ABNORMAL FAULTS OR RESETS. ALSO, THE DEVICE IS OPERATING NORMALLY. THE HIGHER POWER CONSUMPTION IS DUE TO THE HIGHER OUTPUT SETTINGS. TO DATE, THE DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION INDICATED THAT THIS DEVICE WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21108588&pc=LWP | 2124215-2024-83345 | 2024-12-17 05:00 | 2025-01-08 05:00 | Injury | Premature Discharge of Battery; Battery Problem | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITING WAS PREMATURE BATTERY DEPLETION. THE BATTERY HAD DEPLETED BY FOUR YEARS TO THREE YEARS IN APPROXIMATELY FIVE MONTHS. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS CONFIRMED THE BATTERY APPEARED TO BE DEPLETING MORE QUICKLY THAN EXPECTED. DEVICE REPLACEMENT WAS RECOMMENDED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. CURRENTLY, EVIDENCE SUGGESTS THAT THIS PRODUCT REMAINS IN-SERVICEEvent Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITING WAS PREMATURE BATTERY DEPLETION. THE BATTERY HAD DEPLETED BY FOUR YEARS TO THREE YEARS IN APPROXIMATELY FIVE MONTHS. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS CONFIRMED THE BATTERY APPEARED TO BE DEPLETING MORE QUICKLY THAN EXPECTED. DEVICE REPLACEMENT WAS RECOMMENDED. SUBSEQUENTLY THIS PACEMAKER WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVENT PATIENT EFFECT WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20965748&pc=LWP | 2124215-2024-80222 | 2024-12-06 05:00 | 2024-12-18 05:00 | Injury | Premature Discharge of Battery; Battery Problem | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS IMPLANTABLE DEVICE WAS SUSPECTED TO BE DEPLETING PREMATURELY. THIS DEVICE WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS PRODUCT IS BEING EVALUATED IN OUR POST MARKET QUALITY ASSURANCE LABORATORY. THIS REPORT WILL BE UPDATED WHEN EVALUATION IS COMPLETE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20945892&pc=LWP | 2124215-2024-79538 | 2024-12-12 05:00 | 2024-12-16 05:00 | Injury | Premature Discharge of Battery; Battery Problem | ACCOLADE SR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS IMPLANTABLE DEVICE WAS SUSPECTED TO BE DEPLETING PREMATURELY. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS CONFIRMED THE BATTERY APPEARED TO BE DEPLETING MORE QUICKLY THAN EXPECTED. DEVICE REPLACEMENT WAS RECOMMENDED. THIS DEVICE REMAINS IN SERVICE WITHOUT IMPACT TO THERAPY. NO ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS IMPLANTABLE DEVICE WAS SUSPECTED TO BE DEPLETING PREMATURELY. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS CONFIRMED THE BATTERY APPEARED TO BE DEPLETING MORE QUICKLY THAN EXPECTED. DEVICE REPLACEMENT WAS RECOMMENDED. THIS DEVICE REMAINS IN SERVICE WITHOUT IMPACT TO THERAPY. NO ADVERSE PATIENT EFFECTS WERE REPORTED. DEVICE WAS EXPLANTED AND RETURNED TO BOSTON SCIENTIFIC. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20925572&pc=LWP | 2124215-2024-79011 | 2024-09-01 04:00 | 2024-12-13 05:00 | Injury | Premature Discharge of Battery; Battery Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS PACEMAKER WAS SUSPECTED TO BE DEPLETING PREMATURELY. CONSEQUENTLY, THE DEVICE WAS EXPLANTED AND REPLACED. AT THIS TIME, THERE IS NO EVIDENCE TO SUGGEST A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED PRIOR TO THE EXPLANT PROCEDURE TO CONFIRM THE PREMATURE BATTERY DEPLETION OBSERVATION. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THE DEVICE IS EXPECTED TO BE RETURNED FOR ANALYSIS.Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS PACEMAKER WAS SUSPECTED TO BE DEPLETING PREMATURELY. CONSEQUENTLY, THE DEVICE WAS EXPLANTED AND REPLACED. AT THIS TIME, THERE IS NO EVIDENCE TO SUGGEST A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED PRIOR TO THE EXPLANT PROCEDURE TO CONFIRM THE PREMATURE BATTERY DEPLETION OBSERVATION. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THE DEVICE IS EXPECTED TO BE RETURNED FOR ANALYSIS. Manufacturer Narrative: AT THIS TIME, THE PRODUCT HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED, ANALYSIS WILL BE PERFORMED, AND THIS REPORT WILL BE UPDATED AT THAT TIME. Manufacturer Narrative: AT THIS TIME, THE PRODUCT HAS NOT BEEN RETURNED. IF THE PRODUCT IS RETURNED, ANALYSIS WILL BE PERFORMED, AND THIS REPORT WILL BE UPDATED AT THAT TIME.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20930687&pc=LWP | 2124215-2024-79129 | 2024-11-13 05:00 | 2024-12-13 05:00 | Injury | Premature Discharge of Battery; Battery Problem | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER WAS SUSPECTED OF PREMATURE BATTERY DEPLETION (PBD) SINCE BATTERY DROPPED FROM 5 AND A HALF YEARS TO 3 YEARS IN A SPAN OF ONE YEAR. A REQUEST WAS MADE TO HAVE DATA FROM THE DEVICE ANALYZED. DATA ANALYSIS CONFIRMED PBD. DEVICE REPLACEMENT WAS RECOMMENDED OR HAVE THE BATTERY STATUS RE-EVALUATED WITHIN 90 DAYS. SUBSEQUENTLY, THIS DEVICE WAS ELECTRICALLY ABANDONED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20917659&pc=LWP | 2124215-2024-78770 | 2024-12-09 05:00 | 2024-12-12 05:00 | Injury | Premature Discharge of Battery; Battery Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS PACEMAKER WAS SUSPECTED TO BE DEPLETING PREMATURELY. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS CONFIRMED THE BATTERY APPEARED TO BE DEPLETING MORE QUICKLY THAN EXPECTED. THIS DEVICE REMAINED IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS PACEMAKER WAS SUSPECTED TO BE DEPLETING PREMATURELY. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS CONFIRMED THE BATTERY APPEARED TO BE DEPLETING MORE QUICKLY THAN EXPECTED. THIS DEVICE REMAINED IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION WAS RECEIVED. A DEVICE REPLACEMENT INTERVAL WAS PROVIDED. ADDITIONAL INFORMATION WAS RECEIVED. THE DEVICE WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS PACEMAKER WAS SUSPECTED TO BE DEPLETING PREMATURELY. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS CONFIRMED THE BATTERY APPEARED TO BE DEPLETING MORE QUICKLY THAN EXPECTED. THIS DEVICE REMAINED IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION WAS RECEIVED. A DEVICE REPLACEMENT INTERVAL WAS PROVIDED. ADDITIONAL INFORMATION WAS RECEIVED. THE DEVICE WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: D6B: UPDATED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21254052&pc=LWP | 2124215-2025-04134 | 2025-01-05 05:00 | 2025-01-28 05:00 | Injury | Premature Discharge of Battery; Battery Problem; Impedance Problem | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THE PACEMAKER BATTERY APPEARED TO BE DEPLETING PREMATURELY, AS THE ESTIMATED REMAINING LONGEVITY DECREASED ABNORMALLY. UPON REVIEWING DATA FROM THE DEVICE, IT WAS NOTED THAT THE RIGHT ATRIAL (RA) LEAD SHOWED LOW PACING IMPEDANCE; HOWEVER, NO ROOT CAUSE WAS IDENTIFIED. FURTHER ANALYSIS CONFIRMED ACCELERATED BATTERY DEPLETION. DEVICE REPLACEMENT WAS RECOMMENDED, AND THE PACEMAKER WAS SUBSEQUENTLY EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE EFFECTS ON THE PATIENT WERE REPORTED.Event Description: IT WAS REPORTED THAT THE PACEMAKER BATTERY APPEARED TO BE DEPLETING PREMATURELY, AS THE ESTIMATED REMAINING LONGEVITY DECREASED ABNORMALLY. UPON REVIEWING DATA FROM THE DEVICE, IT WAS NOTED THAT THE RIGHT ATRIAL (RA) LEAD SHOWED LOW PACING IMPEDANCE; HOWEVER, NO ROOT CAUSE WAS IDENTIFIED. FURTHER ANALYSIS CONFIRMED ACCELERATED BATTERY DEPLETION. DEVICE REPLACEMENT WAS RECOMMENDED, AND THE PACEMAKER WAS SUBSEQUENTLY EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE EFFECTS ON THE PATIENT WERE REPORTED. Manufacturer Narrative: CORRECTED FIELDS: H1 TYPE OF REPORTABLE EVENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21194181&pc=LWP | 2124215-2025-02484 | 2025-01-04 05:00 | 2025-01-20 05:00 | Injury | Premature Discharge of Battery; Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI EL DR | Blurred Vision | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS PACEMAKER WAS EVALUATED IN A HOSPITAL SETTING WITH SYMPTOMS OF DARK SENSATION IN FRONT OF THEIR EYES. IT WAS THEN CONFIRMED THAT THIS PACEMAKER WAS IN SAFETY MODE. A DEVICE CHANGEOUT PROCEDURE WAS PERFORMED AND IT WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. IT IS EXPECTED TO RECEIVE THIS PRODUCT FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21179672&pc=LWP | 2124215-2025-02081 | 2025-01-02 05:00 | 2025-01-17 05:00 | Injury | Premature Discharge of Battery; Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED A RED ALERT DUE TO THE DEVICE FALLING INTO SAFETY MODE. THE PACEMAKER EXPERIENCED PREMATURE BATTERY DEPLETION (PBD). THE DEVICE WAS EXPLANTED AND REPLACED. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20985152&pc=LWP | 2124215-2024-80932 | 2024-12-09 05:00 | 2024-12-20 05:00 | Injury | Premature Discharge of Battery; Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS FOUND TO BE IN SAFETY MODE, THE PHYSICIAN EXPLANTED AND REPLACED THE DEVICE DUE TO A SUSPECTED PREMATURE BATTERY DEPLETION. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE HAS NOT BEEN RETURNED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20993104&pc=LWP | 2124215-2024-81156 | 2024-12-03 05:00 | 2024-12-20 05:00 | Injury | Premature Discharge of Battery; Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXPLANTED AND REPLACED DUE TO GOING TO SAFETY MODE AND EXHIBITING PREMATURE BATTERY DEPLETION (PBD). THE DEVICE IS EXPECTED TO BE RETURNED FOR ANALYSIS. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20967872&pc=LWP | 2124215-2024-80293 | 2024-12-07 05:00 | 2024-12-18 05:00 | Injury | Premature Discharge of Battery; Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS FOUND TO BE IN SAFETY MODE. TECHNICAL SERVICES (TS) WAS CONSULTED AND DEVICE REPLACEMENT WAS RECOMMENDED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. AT THIS TIME, THIS DEVICE REMAINS IN SERVICE.Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS FOUND TO BE IN SAFETY MODE. TECHNICAL SERVICES (TS) WAS CONSULTED AND DEVICE REPLACEMENT WAS RECOMMENDED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. AT THIS TIME, THIS DEVICE REMAINS IN SERVICE. FURTHER INFORMATION WAS RECEIVED THAT THIS DEVICE WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. IT IS EXPECTED TO RECEIVE THIS DEVICE FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20909480&pc=LWP | 2124215-2024-78506 | 2024-10-15 04:00 | 2024-12-11 05:00 | Injury | Premature Discharge of Battery; Incorrect, Inadequate or Imprecise Result or Readings; Communication or Transmission Problem | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SHOWED DIFFICULTIES FOR INTERROGATION. AFTERWARDS, IT WAS DETERMINED THAT IT HAD TRIGGERED SAFETY MODE. THE PACEMAKER WAS RECOMMENDED TO BE EXPLANTED. IT WAS EXPLANTED WITH AN ADDITIONAL ALLEGATION OF PREMATURE BATTERY DEPLETION (PBD) AT THIS TIME. A NEW DEVICE WAS IMPLANTED AT THE SAME PROCEDURE. THE EXPLANTED PACEMAKER IS EXPECTED TO BE RETURNED FOR ANALYSIS. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THIS PACEMAKER SHOWED DIFFICULTIES FOR INTERROGATION. AFTERWARDS, IT WAS DETERMINED THAT IT HAD TRIGGERED SAFETY MODE. THE PACEMAKER WAS RECOMMENDED TO BE EXPLANTED BUT REMAINS IN SERVICE AT THIS TIME. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21035072&pc=LWP | 2124215-2024-82348 | 2024-12-03 05:00 | 2024-12-30 05:00 | Injury | Signal Artifact/Noise | ACCOLADE DR | Ventricular Fibrillation | N970003 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS IMPLANTABLE PACEMAKER WAS EXTERNALLY RESCUED DURING ELECTROCAUTERY AT CHANGE OUT BECAUSE OF NOISE ON THE SURFACE THAT TURNED OUT TO BE VENTRICULAR FIBRILLATION (VF). BOSTON SCIENTIFIC TECHNICAL SERVICES (TS) PROVIDED AN ONLINE CASE STUDY, SUPPORTING THE POSSIBILITY OF ELECTROCAUTERY AT CHANGE OUT CAUSING VENTRICULAR FIBRILLATION. THIS PACEMAKER WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PRODUCT HAS BEEN RETURNED FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21223783&pc=LWP | 2124215-2025-03257 | 2024-12-26 05:00 | 2025-01-23 05:00 | Injury | Signal Artifact/Noise; Failure to Capture; Failure to Convert Rhythm; Low impedance; Defective Device; High Capture Threshold | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS PACEMAKER VISITED THE HOSPITAL. DURING DEVICE INTERROGATION, THE PACEMAKER WAS FOUND TO EXHIBIT DIFFICULTIES CONVERTING VENTRICULAR FIBRILLATION, LOSS OF CAPTURE (LOC), HIGH PACING THRESHOLDS, NOISE, AND LOW OUT OF RANGE PACING IMPEDANCES ON THE RIGHT VENTRICULAR (RV) LEAD. THE PATIENT WAS HOSPITALIZED, AND X-RAYS WERE TAKEN, HOWEVER THE IMAGES WERE DETERMINED TO BE NORMAL BY THE PHYSICIAN. AN EMERGENT SURGICAL INTERVENTION WAS PERFORMED, THE RV LEAD WAS EXPLANTED AND REPLACED WITH A NEW LEAD. THE PHYSICIAN REPORTED THAT DURING THE PLACEMENT AND CONNECTION OF THE NEW RV LEAD TO THIS PACEMAKER DEVICE, THE IMPEDANCE VALUES OF THE NEW RV LEAD WERE NOTED TO HAVE DROPPED FROM 800 TO 500 OHMS. THE PHYSICIAN SUSPECTED THE CAUSE WAS DUE TO A DAMAGE TO THE INTERNAL CIRCUITS OF THE DEVICE AND OPTED TO REPLACE THE PACEMAKER WITH A NEW DEVICE. THE DEVICE WAS SUCCESSFULLY IMPLANTED AND CONNECTED TO THE NEW RV LEAD. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THE EXPLANTED PACEMAKER AND RV LEAD ARE EXPECTED TO RETURN FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21152264&pc=LWP | 2124215-2025-01327 | 2024-12-19 05:00 | 2025-01-14 05:00 | Injury | Signal Artifact/Noise; High impedance; Over-Sensing; High Capture Threshold | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER WAS EXPLANTED AND REPLACED DUE TO A COMPONENT MALFUNCTION. IT WAS ALSO REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD WAS CAPPED AND LEFT IN-SITU. THE RV LEAD WAS REPLACED. NO ADDITIONAL INFORMATION WAS PROVIDED. OUTSIDE OF THE REVISION, NO ADVERSE PATIENT EFFECTS WERE REPORTED. RECEIVED ADDITIONAL INFORMATION THE PATIENT WAS REFERRED TO THE HOSPITAL DUE TO RV LEAD ISSUES. THE RV LEAD THRESHOLD WAS 6.0V (VOLTS) AT 0.8MS (MILLISECONDS) AND 5.0V AT 1.5MS IN BIPOLAR POLARITY AND IN UNIPOLAR 6.5V AT 0.8MS. THE RV PACING IMPEDANCE WAS 2639 OHMS IN UNIPOLAR AND 2850 OHMS IN BIPOLAR. THE TRENDS SHOWED A STEADY INCREASE FROM 1900 OHMS OVER LAST 12 MONTHS. OVERSENSING FROM NOISE WAS ALSO NOTED. THE PHYSICIAN ELECTED TO REPLACE BOTH THE DEVICE AND THE RV LEAD. THE DEVICE WILL NOT RETURN FOR ANALYSIS. OUTSIDE OF THE REVISION, NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21200811&pc=LWP | 2124215-2025-02700 | 2024-12-31 05:00 | 2025-01-21 05:00 | Injury | Signal Artifact/Noise; Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE PACEMAKER AND RV LEAD EXHIBITED NOISE. UPON INTERROGATION, THE DEVICE WAS FOUND TO HAVE REVERTED TO SAFETY MODE. SUBSEQUENTLY, THE DEVICE AND RV LEAD WERE EXPLANTED AND REPLACED SUCCESSFULLY. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21214635&pc=LWP | 2124215-2025-03067 | 2025-01-07 05:00 | 2025-01-23 05:00 | Injury | Signal Artifact/Noise; Over-Sensing | ACCOLADE MRI DR | Syncope/Fainting | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS PACEMAKER SYSTEM EXPERIENCED A SYNCOPAL EPISODE. TECHNICAL SERVICES (TS) REVIEWED THE EVENT AND IDENTIFIED NOISE OVERSENSING ASSOCIATED WITH THE MINUTE VENTILATION (MV) FEATURE, WHICH HAD BEEN DEACTIVATED BY THE SIGNAL ARTIFACT MONITOR (SAM). ADDITIONALLY, A HIGH NUMBER OF RIGHT VENTRICULAR AUTOMATIC THRESHOLD (RVAT) TESTS WERE OBSERVED. TS PROVIDED REPROGRAMMING OPTIONS AND RECOMMENDED FURTHER EVALUATION OF THE RIGHT VENTRICULAR (RV) LEAD, AS A LEAD ISSUE WAS SUSPECTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21103453&pc=LWP | 2124215-2025-00245 | 2021-04-15 04:00 | 2025-01-08 05:00 | Injury | Signal Artifact/Noise; Over-Sensing | ACCOLADE DR | Device Overstimulation of Tissue | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM EXHIBITED RIGHT ATRIAL (RA) NOISE THAT WAS BEING OVERSENSED WHICH RESULTED IN INAPPROPRIATE STORED ATRIAL TACHY RESPONSE (ATR) EPISODES. TECHNICAL SERVICES REVIEWED AND APPEARED THE NOISE WAS LIKELY RELATED TO THE LEAD BEING PROGRAMMED TOO UNIPOLAR. ISOMETRICS WAS PERFORMED AND THE NOISE COULD BE RE-CREATED. THE NOISE APPEARED TO BE DUE TO MUSCLE NOISE. THE PATIENT DID FEEL LIKE HER DEVICE WAS NOT RIGHT AND WAS HAVING PAUSES OR PALPITATIONS. THE FIELD REPRESENTATIVE DID NOT BELIEVE HER SYMPTOMS WERE RELATED TO THE DEVICE. SUBSEQUENTLY, THE DEVICE WAS EXPLANTED AND REPLACED, AND THE RA LEAD REMAINS IN SERVICE. THE DEVICE HAS BEEN RETURNED AND IS IN THE PROCESS OF DEVICE ANALYSIS. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21261225&pc=LWP | 2124215-2025-04247 | 2025-01-09 05:00 | 2025-01-28 05:00 | Injury | Signal Artifact/Noise; Over-Sensing; Pacing Problem | ACCOLADE MRI EL DR | Bradycardia | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT DURING THE PATIENT'S HOSPITAL STAY FOLLOWING A PACEMAKER GENERATOR REPLACEMENT, THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED NOISE OVERSENSING, RESULTING IN PACING INHIBITION LASTING BETWEEN ONE AND TWO SECONDS. THE PATIENT ALSO EXPERIENCED DIZZINESS ASSOCIATED WITH THIS EVENT. CONSEQUENTLY, A REVISION SURGERY WAS PERFORMED, DURING WHICH AN ADDITIONAL LEAD WAS IMPLANTED, AND THE RV LEAD WAS REPOSITIONED. NO FURTHER ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21035603&pc=LWP | 2124215-2024-82364 | 2024-12-04 05:00 | 2024-12-30 05:00 | Injury | Signal Artifact/Noise; Over-Sensing; Pacing Problem; Connection Problem; High Capture Threshold | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THERE WERE HIGH CAPTURE THRESHOLDS OF THE RIGHT VENTRICULAR (RV) LEAD OF THIS PACEMAKER SYSTEM SHORTLY AFTER INITIAL IMPLANT. THE PHYSICIAN ELECTED TO EXPLANT AND REPLACE THE RV LEAD. DURING THIS PROCEDURE, AFTER POCKET CLOSURE, THERE WAS NOISE ARTIFACT PRESENT ON THE RV LEAD CHANNEL WHICH RESULTED IN OVERSENSING AND PACING INHIBITION. THE PHYSICIAN REOPENED THE POCKET TO VERIFY THE CONNECTIONS WERE SECURE. TECHNICAL SERVICES (TS) NOTED THAT IT WAS MOST LIKELY DUE TO AIR BUBBLES AND SUGGESTED REPROGRAMMING OPTIONS AS TROUBLESHOOTING. THE PHYSICIAN OPTED TO REPLACE THE PACEMAKER WHICH RESOLVED THE ARTIFACT. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THE PACEMAKER HAS BEEN RETURNED TO BOSTON SCIENTIFIC FOR FURTHER INVESTIGATION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21185787&pc=LWP | 2124215-2025-02204 | 2025-01-03 05:00 | 2025-01-17 05:00 | Injury | Signal Artifact/Noise; Over-Sensing; Pacing Problem; Incorrect, Inadequate or Imprecise Result or Readings; Battery Problem | ACCOLADE MRI DR | Dyspnea; Presyncope; Syncope/Fainting | P150012 | 1 | Event Description: IT WAS REPORTED THAT DURING DEVICE INTERROGATION, THE PACEMAKER HAD ENTERED SAFETY MODE. UPON CONFIRMING THE PATIENT WAS PACER DEPENDENT, THE DECISION WAS MADE TO TRANSFER THE PATIENT TO THE HOSPITAL TO PREPARE FOR AN URGENT DEVICE REPLACEMENT PROCEDURE. THE PATIENT PASSED OUT ONCE IN THE DOCTOR'S OFF AND TWICE IN THE AMBULANCE. DURING THE TRANSFER, THE PATIENT ALSO EXPERIENCED PACING INHIBITION, SHORTNESS OF BREATH, AND LIGHTHEADEDNESS. IT WAS SUSPECTED THAT THESE SYMPTOMS WERE ATTRIBUTED TO OVERSENSED NOISE WITH UNIPOLAR PACING. THE PACEMAKER WAS EXPLANTED AND REPLACED THAT SAME DAY, DESPITE PREVIOUS PLANS FOR REPLACEMENT IN FEBRUARY. THE PATIENT FELT MUCH BETTER AFTER THE PROCEDURE, WITH NO FURTHER COMPLAINTS. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. PRODUCT RETURN WAS REQUESTED. Manufacturer Narrative: THE LOCAL AREA FIELD REPRESENTATIVE WAS CONTACTED FOR ADDITIONAL INFORMATION REGARDING PRODUCT RETURN STATUS. AT THIS TIME, NO FURTHER INFORMATION IS AVAILABLE. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THIS REPORT WILL BE UPDATED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21279808&pc=LWP | 2124215-2025-04958 | 2025-01-14 05:00 | 2025-01-30 05:00 | Injury | Unexpected Therapeutic Results | ACCOLADE MRI DR | Bradycardia; Malaise; Loss of consciousness; Syncope/Fainting | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH THIS PACEMAKER SYSTEM PRESENTED TO THE EMERGENCY ROOM (ER) DUE TO A HEART RATE OF 44 BEATS PER MINUTE (BPM), DESPITE DEVICE PROGRAMMING SET TO 60 BPM AND 70 BPM. THE PATIENT WAS KEPT IN THE ER FOR HALF A DAY BEFORE BEING TRANSFERRED TO ANOTHER HOSPITAL. A FIELD REPRESENTATIVE CHECKED THE DEVICE AT THE SECONDARY HOSPITAL, AND IT WAS OPERATING AS EXPECTED. A FEW DAYS LATER, THE PATIENT'S HEART RATE WAS 54 BPM AND HE HAD DIFFICULTY STAYING AWAKE. THE PATIENT WAS FEELING UNWELL AND HAD COLLAPSED ONTO A CHAIR. THE FIELD REPRESENTATIVE AND THE PHYSICIAN/CLINIC WILL DISCUSS FURTHER ACTIONS. THIS PACEMAKER SYSTEM REMAINS IN SERVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21220269&pc=LWP | MW5165150 | 2025-01-13 05:00 | 2025-01-22 05:00 | Malfunction |  | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: DEVICE REMOVED FOR ACCELERATED BATTERY DEPLETION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21196662&pc=LWP | 21196662 | 2024-12-16 05:00 | 2025-01-21 05:00 | Malfunction |  | ACCOLADE¿ EL DR | Arrhythmia; No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: PACEMAKER UNDER BOSTON SCIENTIFIC ADVISORY FOR HIGH BATTERY IMPEDANCE MAY INITIATE SAFETY MODE. RECOMMENDATION FOR PATIENTS WITH 4 YEARS OR LESS BATTERY LONGEVITY TO HAVE PROMPT GENERATOR CHANGE. PATIENT HAD ~2 YEARS BATTERY LONGEVITY, UNDERLYING RHYTHM JUNCTIONAL--HIGH PROBABILITY TO BE SYMPTOMATIC IF SAFETY MODE OCCURS. PACEMAKER GENERATOR CHANGE DONE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21196665&pc=LWP | 21196665 | 2024-12-16 05:00 | 2025-01-21 05:00 | Malfunction |  | ACCOLADE¿ MRI EL DR | No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: PACEMAKER UNDER BOSTON SCIENTIFIC ADVISORY FOR HIGH BATTERY IMPEDANCE MAY INITIATE SAFETY MODE. RECOMMENDATION FOR PATIENTS DEPENDENT ON PACEMAKER AND/OR 4 YEARS OR LESS BATTERY LONGEVITY TO HAVE PROMPT GENERATOR CHANGE. PATIENT HAD ~4.5 YEARS BATTERY LONGEVITY, DEPENDENT ON PACEMAKER. PACEMAKER GENERATOR CHANGE DONE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21196729&pc=LWP | 21196729 | 2024-12-16 05:00 | 2025-01-21 05:00 | Malfunction |  | ACCOLADE¿ EL DR | No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: PACEMAKER UNDER BOSTON SCIENTIFIC ADVISORY FOR HIGH BATTERY IMPEDANCE MAY INITIATE SAFETY MODE. RECOMMENDATION FOR PATIENTS DEPENDENT ON PACEMAKER AND/OR 4 YEARS OR LESS BATTERY LONGEVITY TO HAVE PROMPT GENERATOR CHANGE. PATIENT HAD ~4.5 YEARS BATTERY LONGEVITY, DEPENDENT ON PACEMAKER. PACEMAKER GENERATOR CHANGE DONE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21178307&pc=LWP | 21178307 | 2024-12-16 05:00 | 2025-01-17 05:00 | Malfunction |  | ACCOLADE¿ EL DR | Ventilator Dependent |  | 1 | Event Description: BSC ADVISORY FOR HIGH BATTERY IMPEDANCE MAY INITIATE SAFETY MODE. PATIENT IS DEPENDENT ON DEVICE WITH ~2 YEARS BATTERY REMAINING. RECOMMENDATION IS TO PROPHYLACTICALLY CHANGE PACEMAKER IF BATTERY </= 4 YEARS REMAINING. PACEMAKER GENERATOR CHANGED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21254045&pc=LWP | 2124215-2025-04121 | 2024-12-31 05:00 | 2025-01-28 05:00 | Malfunction | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | Fall | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS IMPLANTABLE PACEMAKER DEVICE FELL TO THE GROUND AND STATED THAT DEVICE WAS NOT FUNCTIONING. AS OF THIS TIME, THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21136749&pc=LWP | 2124215-2025-00973 | 2024-12-18 05:00 | 2025-01-13 05:00 | Malfunction | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI SR | Unspecified Heart Problem | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS IMPLANTABLE PACEMAKER INDICATED THAT HIS HEART WAS NOT DOING GOOD AND DESCRIBED IT AS SEVERE MISHAP. BOSTON SCIENTIFIC TECHNICAL SERVICES (TS) RECOMMENDED FOLLOWING UP WITH HIS CLINIC. THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21013058&pc=LWP | 2124215-2024-81792 | 2024-11-26 05:00 | 2024-12-24 05:00 | Malfunction | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER DEVICE HAS TRIGGERED END OF LIFE (END OF LIFE (EOL). AS OF THIS TIME, THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20957138&pc=LWP | 2124215-2024-80039 | 2024-10-02 04:00 | 2024-12-17 05:00 | Malfunction | Adverse Event Without Identified Device or Use Problem | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS IMPLANTABLE PACEMAKER STATED HAVING SOME ISSUES AFTER REPROGRAMMING THE DEVICE. THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS IMPLANTABLE PACEMAKER STATED HAVING SOME ISSUES AFTER REPROGRAMMING THE DEVICE. THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION RECEIVED INDICATES THAT THE DEVICE WAS WORKING AS PROGRAMMED AND THE PATIENT JUST DID NOT LIKE THE RATE RESPONSE. Manufacturer Narrative: THIS SUPPLEMENTAL REPORT WAS CREATED TO CAPTURE THE ADDITIONAL INFORMATION RECEIVED THAT THE DEVICE WAS WORKING AS PROGRAMMED AND THE PATIENT JUST DID NOT LIKE THE RATE RESPONSE AND A CORRECTION TO THE INITIAL EMDR SINCE THIS NO LONGER MEET REPORTABLE CRITERIA.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21213950&pc=LWP | 2124215-2025-03010 | 2025-01-14 05:00 | 2025-01-22 05:00 | Malfunction | Battery Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS IMPLANTABLE DEVICE WAS SUSPECTED TO BE DEPLETING PREMATURELY DUE TO A DROP IN ESTIMATED LONGEVITY OF THREE YEARS OVER THE COURSE OF ONE YEAR. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. NO DATA ANALYSIS HAS BEEN COMPLETED AT THIS TIME. THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21214415&pc=LWP | 2124215-2025-03089 | 2024-12-26 05:00 | 2025-01-22 05:00 | Malfunction | Battery Problem | ACCOLADE EL DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS FIELD REPRESENTATIVE IS CONCERNED ABOUT THE BATTERY OF THIS IMPLANTABLE PACEMAKER. THIS DEVICE HAD 2.5 YEARS REMAINING AFTER 9 YEARS FROM THE IMPLANT DATE. THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21163610&pc=LWP | 2124215-2025-01632 | 2024-12-26 05:00 | 2025-01-15 05:00 | Malfunction | Battery Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THERE WERE CONCERNS AS THE PACEMAKER EXPERIENCED A SIGNIFICANT DECREASE IN BATTERY LONGEVITY. TECHNICAL SERVICES (TS) REVIEWED THE REPORTS AND NOTED THAT THE LONGEVITY DECREASE DOES NOT APPEAR TO BE APPROPRIATE DEVICE BEHAVIOR. TS REQUESTED THE CALLER TO ENSURE THE CORRECT DATA WAS SENT. TS STATED THAT THE PATIENT ALSO EXPERIENCED SOME ATRIAL ARRHYTHMIAS THAT COULD HAVE CONTRIBUTED TO THE LONGEVITY DECREASE. THE DEVICE REMAINS IN USE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21054910&pc=LWP | 2124215-2024-82795 | 2024-12-04 05:00 | 2025-01-02 05:00 | Malfunction | Battery Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER DEVICE HAS LESS THAN THREE MONTHS BATTERY REMAINING, HOWEVER THE DEVICE WAS NOT YET REACHING ELECTIVE REPLACEMENT INDICATOR (ERI). AS OF THIS TIME, THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21050631&pc=LWP | 2124215-2024-82718 | 2024-12-02 05:00 | 2024-12-31 05:00 | Malfunction | Battery Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER HAS A BATTERY ANOMALY AND NEEDED A MONITORING WITH THE COMMUNICATOR. THE DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20951718&pc=LWP | 2124215-2024-79761 | 2023-12-05 05:00 | 2024-12-17 05:00 | Malfunction | Battery Problem | ACCOLADE EL DR | No Clinical Signs, Symptoms or Conditions | N970003/S167 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER LONGEVITY DROPPED FROM 2.5 YEARS TO 4 MONTHS IN A SPAN OF 1 YEAR AND 4 MONTHS. ALSO, PATIENT HAD CHRONIC ATRIAL FIBRILLATION (AF). BOSTON SCIENTIFIC TECHNICAL SERVICES (TS) DISCUSSED THAT DEVICE NEED FILES FOR ANALYSIS. DATA ANALYSIS SHOWED THAT THE DEVICE IS NOT DEPLETING PREMATURELY AND THE DROP IN LONGEVITY IS DUE TO THE INCREASE IN POWER CONSUMPTION CAUSED BY HIGH-RATE ATRIAL SENSING. TO DATE, THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION INDICATED THAT THIS DEVICE WAS EXPLANTED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY, THIS DEVICE WAS THOROUGHLY INSPECTED AND ANALYZED. REVIEW OF DEVICE MEMORY CONFIRMED THAT A LOW VOLTAGE ALERTCODE 1003) WAS NOT RECORDED. A HIGH CURRENT DRAIN WAS DETECTED, BUT THE BATTERY STILL SUPPORTED FULL DEVICE FUNCTION. DETAILED ANALYSIS CONFIRMED THE IDENTIFIED HIGH CURRENT DRAIN WAS A RESULT OF COMPROMISED CAPACITORS. THIS RESULTED IN THE OBSERVED PREMATURE BATTERY DEPLETION. IT WAS DETERMINED THAT THE CAPACITORS WERE COMPROMISED DUE TO THE PRESENCE OF EXCESS HYDROGEN IN THE DEVICE CASE. BOSTON SCIENTIFIC ISSUED A FIELD SAFETY NOTICE REGARDING A SUBSET OF DEVICES IN THE ACCOLADE PACEMAKER FAMILY THAT HAVE A POTENTIAL OF EXHIBITING THIS BEHAVIOR. THIS PARTICULAR DEVICE IS INCLUDED IN THE HYDROGEN INDUCED PREMATURE DEPLETION ADVISORY POPULATION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21279777&pc=LWP | 2124215-2025-04912 | 2022-08-31 04:00 | 2025-01-30 05:00 | Malfunction | Battery Problem; Data Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXHIBITING DATA ISSUES. THE PACEMAKER POWER CONSUMPTION INCREASED DUE TO ATRIAL FIBRILLATION (AF). TECHNICAL SERVICES (TS) WAS CONSULTED AND EXPLAINED BATTERY WILL PORTRAIT ACCURATE LONGEVITY AFTERWARDS. THE DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20920043&pc=LWP | 2124215-2024-78835 | 2024-11-15 05:00 | 2024-12-12 05:00 | Malfunction | Battery Problem; Interrogation Problem | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER COULD NOT BE INTERROGATED, DESPITE ATTEMPTS BY THREE DIFFERENT HOSPITALS. THERE WAS ALSO NO MAGNET RESPONSE. THE REMAINING LONGEVITY IN (B)(6) 2024 HAD BEEN TWO YEARS. FOLLOW-UP WITH THE FIELD YIELDED NO FURTHER INFORMATION. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21217487&pc=LWP | 2124215-2025-03117 | 2024-12-31 05:00 | 2025-01-23 05:00 | Malfunction | Connection Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT, DURING AN INITIAL IMPLANT PROCEDURE, THIS PACEMAKER EXHIBITED CONNECTION ISSUES. THE PROGRAMMER DISPLAYED A TELEMETRY WAND OUT OF RANGE MESSAGE WHEN THE WAND WAS REMOVED, AS THE WAND WAS NECESSARY FOR COMMUNICATION DUE TO THE PROGRAMMED SETTINGS. THIS PACEMAKER WAS SUCCESSFULLY REPLACED, AND NEW DEVICE OF THE SAME MODEL WAS IMPLANTED INSTEAD. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21045389&pc=LWP | 2124215-2024-82552 | 2024-12-04 05:00 | 2024-12-31 05:00 | Malfunction | Connection Problem | ACCOLADE DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS RIGHT VENTRICULAR (RV) LEAD SET SCREW WAS NOT OPERATING NORMALLY AND THE TORQUE WRENCH WOULD JUST SPIN THE SET SCREW IN THE SOCKET. THEY REMAINED WITH THE PROCEDURE WITHOUT CUTTING THE HEADER OR LEAD. THE LEAD REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THIS RIGHT VENTRICULAR (RV) LEAD SET SCREW WAS NOT OPERATING NORMALLY AND THE TORQUE WRENCH WOULD JUST SPIN THE SET SCREW IN THE SOCKET. THEY REMAINED WITH THE PROCEDURE WITHOUT CUTTING THE HEADER OR LEAD. THE LEAD REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THE PHYSICIAN WAS ABLE TO USE RECOMMENDED TECHNIQUES TO FREE THE LEAD FROM THE HEADER. THE DEVICE WHICH WAS EXPLANTED IN THE PROCEDURE DUE TO NORMAL BATTERY DEPLETION IS EXPECTED TO BE RETURNED FOR ANALYSIS. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21285884&pc=LWP | 2124215-2025-05091 | 2021-06-07 04:00 | 2025-01-31 05:00 | Malfunction | Defective Device | ACCOLADE DR | No Clinical Signs, Symptoms or Conditions | N970003/S167 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER HAD ARTIFACT ON TRACINGS. BOSTON SCIENTIFIC TECHNICAL SERVICES (TS) DISCUSSED THAT IT IS BLANKED. THIS DEVICE WAS EXPLANTED AND REPLACED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY, THIS DEVICE WAS THOROUGHLY INSPECTED AND ANALYZED. A HIGH CURRENT DRAIN WAS DETECTED, BUT THE BATTERY STILL SUPPORTED FULL DEVICE FUNCTION. DETAILED ANALYSIS CONFIRMED THE IDENTIFIED HIGH CURRENT DRAIN WAS A RESULT OF COMPROMISED CAPACITORS. THIS RESULTED IN THE OBSERVED PREMATURE BATTERY DEPLETION. IT WAS DETERMINED THAT THE CAPACITORS WERE COMPROMISED DUE TO THE PRESENCE OF EXCESS HYDROGEN IN THE DEVICE CASE. BOSTON SCIENTIFIC ISSUED A FIELD SAFETY NOTICE REGARDING A SUBSET OF DEVICES IN THE ACCOLADE PACEMAKER FAMILY THAT HAVE A POTENTIAL OF EXHIBITING THIS BEHAVIOR. THIS PARTICULAR DEVICE IS INCLUDED IN THE HYDROGEN INDUCED PREMATURE DEPLETION ADVISORY POPULATION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21192545&pc=LWP | 2124215-2025-02439 | 2024-12-26 05:00 | 2025-01-20 05:00 | Malfunction | Defective Device | ACCOLADE MRI DR | Bradycardia | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER DEVICE WAS NOT FUNCTIONING PROPERLY AS THE PATIENT EXPERIENCED A LOW HEART RATE OF THIRTIES. BOSTON SCIENTIFIC TECHNICAL SERVICES (TS) REFERRED THE PATIENT TO THE CLINIC FOR A REPORT AND PATIENT SYMPTOMS. AS OF THIS TIME, THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21161770&pc=LWP | 2124215-2025-01567 | 2024-12-19 05:00 | 2025-01-15 05:00 | Malfunction | Defective Device | ACCOLADE MRI DR | Ischemia Stroke | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER DEVICE WAS NOT FUNCTIONING PROPERLY AND PATIENT EXPERIENCED STROKE LIKED SYMPTOMS. BOSTON SCIENTIFIC TECHNICAL SERVICES (TS) STATED THAT THERE WAS NO HISTORY OF DEVICE ISSUES. AS OF THIS TIME, THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21045272&pc=LWP | 2124215-2024-82529 | 2024-07-26 04:00 | 2024-12-31 05:00 | Malfunction | Defective Device | ACCOLADE EL DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER WAS RECEIVED BY POST MARKET QUALITY ASSURANCE WITHOUT A KNOWN REASON FOR EXPLANT. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY, THIS DEVICE WAS THOROUGHLY INSPECTED AND ANALYZED. A HIGH CURRENT DRAIN WAS DETECTED, BUT THE BATTERY STILL SUPPORTED FULL DEVICE FUNCTION. DETAILED ANALYSIS CONFIRMED THE IDENTIFIED HIGH CURRENT DRAIN WAS A RESULT OF COMPROMISED CAPACITORS. THIS RESULTED IN THE OBSERVED PREMATURE BATTERY DEPLETION. IT WAS DETERMINED THAT THE CAPACITORS WERE COMPROMISED DUE TO THE PRESENCE OF EXCESS HYDROGEN IN THE DEVICE CASE. BOSTON SCIENTIFIC ISSUED A FIELD SAFETY NOTICE REGARDING A SUBSET OF DEVICES IN THE ACCOLADE PACEMAKER FAMILY THAT HAVE A POTENTIAL OF EXHIBITING THIS BEHAVIOR. THIS PARTICULAR DEVICE IS INCLUDED IN THE HYDROGEN INDUCED PREMATURE DEPLETION ADVISORY POPULATION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21027915&pc=LWP | 2124215-2024-82151 | 2024-11-27 05:00 | 2024-12-27 05:00 | Malfunction | Defective Device | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS IMPLANTABLE PACEMAKER DEVICE RECEIVED A CALL FROM THE DEVICE CLINIC THAT THE PACEMAKER NEEDED TO BE CHECKED AND REPROGRAMMED. THE DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21012241&pc=LWP | 2124215-2024-81758 | 2024-07-16 04:00 | 2024-12-24 05:00 | Malfunction | Defective Device | ACCOLADE DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS PRODUCT WAS RETURNED WITH NO REPORTED PRODUCT PERFORMANCE ISSUES OR ADVERSE PATIENT EFFECTS. ANALYSIS COMPLETED IN OUR POST MARKET QUALITY ASSURANCE LABORATORY DETERMINED THAT THE DEVICE WAS OPERATING IN SAFETY MODE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY, THIS DEVICE WAS THOROUGHLY INSPECTED AND ANALYZED. DEVICE TELEMETRY DATA CONFIRMED IT WAS OPERATING IN SAFETY MODE/DETERMINED IT HAD UNDERGONE RESETS, AND THAT CRITICAL THERAPY REMAINED AVAILABLE. ENGINEERS DETERMINED THAT THIS DEVICE WAS DEMONSTRATING BEHAVIOR CONSISTENT WITH A HIGH INTERNAL CELL IMPEDANCE WITHIN THE BATTERY. THE HIGH INTERNAL IMPEDANCE PUT THIS DEVICE AT RISK OF EXPERIENCING TRANSIENT VOLTAGE DECREASES (AND ASSOCIATED RESETS) DURING PERIODS OF HIGH-POWER CONSUMPTION. WHEN BATTERY VOLTAGE DROPS BELOW A MINIMUM THRESHOLD, A SYSTEM RESET IS PERFORMED. IF THREE SYSTEM RESETS OCCUR WITHIN A 48-HOUR PERIOD, THE DEVICE IS DESIGNED TO IMMEDIATELY ENTER SAFETY MODE OPERATION TO MAINTAIN BACK-UP PACING WITH PRE-DEFINED, NON-PROGRAMMABLE SETTINGS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20996894&pc=LWP | 2124215-2024-81275 | 2024-07-16 04:00 | 2024-12-21 05:00 | Malfunction | Defective Device | ACCOLADE EL DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER DEVICE WAS RECEIVED BY POST MARKET QUALITY ASSURANCE WITHOUT A KNOWN REASON FOR EXPLANT. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY, THIS DEVICE WAS THOROUGHLY INSPECTED AND ANALYZED. A HIGH CURRENT DRAIN WAS DETECTED, BUT THE BATTERY STILL SUPPORTED FULL DEVICE FUNCTION. DETAILED ANALYSIS CONFIRMED THE IDENTIFIED HIGH CURRENT DRAIN WAS A RESULT OF COMPROMISED CAPACITORS. THIS RESULTED IN THE OBSERVED PREMATURE BATTERY DEPLETION. IT WAS DETERMINED THAT THE CAPACITORS WERE COMPROMISED DUE TO THE PRESENCE OF EXCESS HYDROGEN IN THE DEVICE CASE. BOSTON SCIENTIFIC ISSUED A FIELD SAFETY NOTICE REGARDING A SUBSET OF DEVICES IN THE ACCOLADE PACEMAKER FAMILY THAT HAVE A POTENTIAL OF EXHIBITING THIS BEHAVIOR. THIS PARTICULAR DEVICE IS INCLUDED IN THE HYDROGEN INDUCED PREMATURE DEPLETION ADVISORY POPULATION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20996895&pc=LWP | 2124215-2024-81276 | 2024-07-16 04:00 | 2024-12-21 05:00 | Malfunction | Defective Device | ACCOLADE DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER WAS RECEIVED BY POST MARKET QUALITY ASSURANCE WITHOUT A KNOWN REASON FOR EXPLANT. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20977304&pc=LWP | 2124215-2024-80590 | 2024-07-22 04:00 | 2024-12-19 05:00 | Malfunction | Defective Device | ACCOLADE SR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER WAS RECEIVED BY POST MARKET QUALITY ASSURANCE WITHOUT A KNOWN REASON FOR EXPLANT. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY, THIS DEVICE WAS THOROUGHLY INSPECTED AND ANALYZED. A HIGH CURRENT DRAIN WAS DETECTED, BUT THE BATTERY STILL SUPPORTED FULL DEVICE FUNCTION. DETAILED ANALYSIS CONFIRMED THE IDENTIFIED HIGH CURRENT DRAIN WAS A RESULT OF COMPROMISED CAPACITORS. THIS RESULTED IN THE OBSERVED PREMATURE BATTERY DEPLETION. IT WAS DETERMINED THAT THE CAPACITORS WERE COMPROMISED DUE TO THE PRESENCE OF EXCESS HYDROGEN IN THE DEVICE CASE. BOSTON SCIENTIFIC ISSUED A FIELD SAFETY NOTICE REGARDING A SUBSET OF DEVICES IN THE ACCOLADE PACEMAKER FAMILY THAT HAVE A POTENTIAL OF EXHIBITING THIS BEHAVIOR. THIS PARTICULAR DEVICE IS INCLUDED IN THE HYDROGEN INDUCED PREMATURE DEPLETION ADVISORY POPULATION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20949287&pc=LWP | 2124215-2024-79702 | 2024-07-09 04:00 | 2024-12-17 05:00 | Malfunction | Defective Device | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER WAS RECEIVED BY POST MARKET QUALITY ASSURANCE WITHOUT A KNOWN REASON FOR EXPLANT. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY, THIS DEVICE WAS THOROUGHLY INSPECTED AND ANALYZED. A HIGH CURRENT DRAIN WAS DETECTED, BUT THE BATTERY STILL SUPPORTED FULL DEVICE FUNCTION. DETAILED ANALYSIS CONFIRMED THE IDENTIFIED HIGH CURRENT DRAIN WAS A RESULT OF COMPROMISED CAPACITORS. THIS RESULTED IN THE OBSERVED PREMATURE BATTERY DEPLETION. IT WAS DETERMINED THAT THE CAPACITORS WERE COMPROMISED DUE TO THE PRESENCE OF EXCESS HYDROGEN IN THE DEVICE CASE. BOSTON SCIENTIFIC ISSUED A FIELD SAFETY NOTICE REGARDING A SUBSET OF DEVICES IN THE ACCOLADE PACEMAKER FAMILY THAT HAVE A POTENTIAL OF EXHIBITING THIS BEHAVIOR. THIS PARTICULAR DEVICE IS INCLUDED IN THE HYDROGEN INDUCED PREMATURE DEPLETION ADVISORY POPULATION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20953454&pc=LWP | 2124215-2024-79835 | 2024-07-12 04:00 | 2024-12-17 05:00 | Malfunction | Defective Device | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: THIS PRODUCT WAS RETURNED TO THE MANUFACTURER WITHOUT ANY REPORT OF PERFORMANCE ISSUES OR ADVERSE PATIENT EFFECTS. THE RETURNED DEVICE WAS ANALYZED AND INITIAL INVESTIGATION NOTED A FAILED LONGEVITY CALCULATION. DETAILED ANALYSIS CONFIRMED PREMATURE BATTERY DEPLETION DUE TO HIGH HYDROGEN INDUCED LEAKY CAPACITORS. Manufacturer Narrative: UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY, THIS DEVICE WAS THOROUGHLY INSPECTED AND ANALYZED. A HIGH CURRENT DRAIN WAS DETECTED, BUT THE BATTERY STILL SUPPORTED FULL DEVICE FUNCTION. DETAILED ANALYSIS CONFIRMED THE IDENTIFIED HIGH CURRENT DRAIN WAS A RESULT OF COMPROMISED CAPACITORS. THIS RESULTED IN THE OBSERVED PREMATURE BATTERY DEPLETION. IT WAS DETERMINED THAT THE CAPACITORS WERE COMPROMISED DUE TO THE PRESENCE OF EXCESS HYDROGEN IN THE DEVICE CASE. BOSTON SCIENTIFIC ISSUED A FIELD SAFETY NOTICE REGARDING A SUBSET OF DEVICES IN THE ACCOLADE PACEMAKER FAMILY THAT HAVE A POTENTIAL OF EXHIBITING THIS BEHAVIOR. THIS PARTICULAR DEVICE IS INCLUDED IN THE HYDROGEN INDUCED PREMATURE DEPLETION ADVISORY POPULATION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20935476&pc=LWP | 2124215-2024-79260 | 2024-07-16 04:00 | 2024-12-15 05:00 | Malfunction | Defective Device | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER WAS RECEIVED BY POST MARKET QUALITY ASSURANCE WITHOUT A KNOWN REASON FOR EXPLANT. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY, THIS DEVICE WAS THOROUGHLY INSPECTED AND ANALYZED. A HIGH CURRENT DRAIN WAS DETECTED, BUT THE BATTERY STILL SUPPORTED FULL DEVICE FUNCTION. DETAILED ANALYSIS CONFIRMED THE IDENTIFIED HIGH CURRENT DRAIN WAS A RESULT OF COMPROMISED CAPACITORS. THIS RESULTED IN THE OBSERVED PREMATURE BATTERY DEPLETION. IT WAS DETERMINED THAT THE CAPACITORS WERE COMPROMISED DUE TO THE PRESENCE OF EXCESS HYDROGEN IN THE DEVICE CASE. BOSTON SCIENTIFIC ISSUED A FIELD SAFETY NOTICE REGARDING A SUBSET OF DEVICES IN THE ACCOLADE PACEMAKER FAMILY THAT HAVE A POTENTIAL OF EXHIBITING THIS BEHAVIOR. THIS PARTICULAR DEVICE IS INCLUDED IN THE HYDROGEN INDUCED PREMATURE DEPLETION ADVISORY POPULATION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21123707&pc=LWP | 2124215-2025-00709 | 2024-12-12 05:00 | 2025-01-10 05:00 | Malfunction | Device Dislodged or Dislocated | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS IMPLANTABLE PACEMAKER DEVICE EXPERIENCED A POKING SENSATION ON THE PATIENT'S SIDE. ADDITIONALLY, THE PATIENT STATED THAT THE LEAD DOESN'T CONNECT, AND THE PATIENT WAS TO HAVE A NEW LEAD IMPLANTED BUT WAS BUMPED FOR SOMEONE WITH AN EMERGENCY. THIS LEAD REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21123693&pc=LWP | 2124215-2025-00635 | 2024-12-22 05:00 | 2025-01-10 05:00 | Malfunction | Failure to Capture | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM EXHIBITED INTERMITTENT LOSS OF CAPTURE (LOC). THE PHYSICIAN OPTED TO PERFORM PROGRAMMING ENHANCEMENTS. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21129691&pc=LWP | 2124215-2025-00789 | 2024-12-11 05:00 | 2025-01-10 05:00 | Malfunction | Failure to Capture | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) WAS EXHIBITING LOSS OF CAPTURE (LOC). TECHNICAL SERVICES (TS) WAS CONSULTED AND RECOMMENDED FURTHER EVALUATION. THE DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20995373&pc=LWP | 2124215-2024-81216 | 2024-12-10 05:00 | 2024-12-20 05:00 | Malfunction | Failure to Capture | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER AND RIGHT VENTRICULAR (RV) LEAD EXHIBITED LOSS OF CAPTURE. THE HEALTH CARE PROFESSIONAL (HCP) INDICATED THAT THE PATIENT WAS HAVING A LOT OF PREMATURE VENTRICULAR CONTRACTIONS (PVC) HOWEVER THE PACING THRESHOLDS HAVE BEEN NORMAL. THE HCP PLANS TO CONTINUE TO MONITOR THIS PATIENT. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20984582&pc=LWP | 2124215-2024-80885 | 2024-12-10 05:00 | 2024-12-19 05:00 | Malfunction | Failure to Capture | ACCOLADE MRI EL DR | Bradycardia; Asystole | P150012 | 1 | Event Description: IT WAS REPORTED THAT LOSS OF CAPTURE WAS SEEN ON THE PACEMAKER ELECTROGRAM (EGM), RESULTING IN HEARTRATES IN THE 30 BEATS-PER-MINUTE RANGE. TECHNICAL SERVICES (TS) WAS CONTACTED, AND TS DISCUSSED PROGRAMMING OPTIONS. THE PACEMAKER SYSTEM REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20947969&pc=LWP | 2124215-2024-79626 | 2024-11-27 05:00 | 2024-12-16 05:00 | Malfunction | Failure to Capture; Device Sensing Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS RECENTLY IMPLANTED PACEMAKER EXHIBITED INTERMITTENT LOSS OF CAPTURE (LOC). TECHNICAL SERVICES (TS) WAS CONSULTED, AND IT WAS FOUND THAT THE PACING THRESHOLD WAS DOWN. TS RECOMMENDED FURTHER TESTING. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.Event Description: IT WAS REPORTED THAT THIS RECENTLY IMPLANTED PACEMAKER EXHIBITED INTERMITTENT LOSS OF CAPTURE (LOC). TECHNICAL SERVICES (TS) WAS CONSULTED, AND IT WAS FOUND THAT THE PACING THRESHOLD WAS DOWN. TS RECOMMENDED FURTHER TESTING. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE. ADDITIONAL INFORMATION WAS RECEIVED, THE PATIENT WAS EVALUATED AT THE CLINIC, THE THRESHOLD WAS STABLE. REPROGRAMMING WAS PERFORMED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20994632&pc=LWP | 2124215-2024-81230 | 2024-12-16 05:00 | 2024-12-20 05:00 | Malfunction | Failure to Capture; Failure to Read Input Signal | ACCOLADE MRI EL DR | Fatigue | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER AND NON-BOSTON SCIENTIFIC RV LEAD WERE EXHIBITING AN INTERMITTENT LOSS OF CAPTURE (LOC). IT WAS NOTED THAT EVERY OTHER VENTRICULAR PACE EVENT EXHIBITED DIFFERENT MORPHOLOGY AND THE PATIENT EXPERIENCED FATIGUE. TECHNICAL SERVICES (TS) PAGED THE LOCAL BOSTON SCIENTIFIC REPRESENTATIVE TO EVALUATE THE DEVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE.Event Description: IT WAS REPORTED THAT THIS PACEMAKER AND NON-BOSTON SCIENTIFIC RV LEAD WERE EXHIBITING AN INTERMITTENT LOSS OF CAPTURE (LOC). IT WAS NOTED THAT EVERY OTHER VENTRICULAR PACE EVENT EXHIBITED DIFFERENT MORPHOLOGY AND THE PATIENT EXPERIENCED FATIGUE. TECHNICAL SERVICES (TS) PAGED THE LOCAL BOSTON SCIENTIFIC REPRESENTATIVE TO EVALUATE THE DEVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE. ADDITIONAL INFORMATION WAS RECEIVED THAT INDICATED THAT THE LOC WAS CONFIRMED AND REPROGRAMMING WAS PERFORMED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21234917&pc=LWP | 2124215-2025-03569 | 2024-12-22 05:00 | 2025-01-24 05:00 | Malfunction | Failure to Capture; Failure to Read Input Signal; Under-Sensing; High Capture Threshold | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS RIGHT VENTRICULAR (RV) LEAD INVOLVED IN THE PACEMAKER SYSTEM EXHIBITED LOSS OF CAPTURE (LOC). THERE WAS FUNCTIONAL UNDERSENSING NOTED AS WELL. THE RV THRESHOLD INCREASED TO 2.5V AND THE R-WAVES DROPPED IN AMPLITUDE. TECHNICAL SERVICES RECOMMENDED VERIFYING THE LEAD LOCATION USING XRAY, AS THIS LEAD HAD RECENTLY BEEN PLACED. THIS RV LEAD AND DEVICE REMAIN IN SERVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21275159&pc=LWP | 2124215-2025-04757 | 2025-01-21 05:00 | 2025-01-30 05:00 | Malfunction | Failure to Capture; High Capture Threshold | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER DEVICE EXHIBITED LOSS OF CAPTURE (LOC) AND HIGH THRESHOLD VALUES. DURING THE FOLLOW-UP VISIT, LEAD TESTING WAS PERFORMED, AND THE RIGHT VENTRICULAR (RV) LEAD THRESHOLDS HAD INCREASED TO 6V. TROUBLESHOOTING WAS PERFORMED TO CONFIRM LOC. ALL OTHER MEASUREMENTS WERE WITHIN THE NORMAL RANGE. THE PHYSICIAN PLANS TO HAVE THIS DEVICE AND LEAD REPLACES SOON. THIS DEVICE CURRENTLY REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21125825&pc=LWP | 2124215-2025-00793 | 2024-10-03 04:00 | 2025-01-10 05:00 | Malfunction | Failure to Capture; High Capture Threshold | ACCOLADE MRI EL DR | Pain | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED LOSS OF CAPTURE (LOC) AND HIGH CAPTURE THRESHOLDS ON THE RIGHT VENTRICULAR (RV) LEAD. TECHNICAL SERVICES (TS) WAS CONSULTED AND RECOMMENDED INCREASING OUTPUTS. THE DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21013040&pc=LWP | 2124215-2024-81776 | 2024-12-16 05:00 | 2024-12-24 05:00 | Malfunction | Failure to Capture; High Capture Threshold | ACCOLADE MRI EL DR | Presyncope | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM WITH A NON-BOSTON SCIENTIFIC RIGHT VENTRICULAR (RV) LEAD EXHIBITED LOSS OF CAPTURE (LOC) AND HIGH RV PACING THRESHOLDS WERE OBSERVED. AN ALERT WAS DETECTED FOR RIGHT VENTRICULAR AUTOMATIC THRESHOLD (RVAT) TEST, INDICATING THRESHOLDS HIGHER THAN EXPECTED OR A SUSPENSION. THE PATIENT HAD A PRE-SYNCOPAL EPISODE. FURTHER EVALUATION WAS RECOMMENDED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM WITH A NON-BOSTON SCIENTIFIC RIGHT VENTRICULAR (RV) LEAD EXHIBITED LOSS OF CAPTURE (LOC) AND HIGH RV PACING THRESHOLDS WERE OBSERVED. AN ALERT WAS DETECTED FOR RIGHT VENTRICULAR AUTOMATIC THRESHOLD (RVAT) TEST, INDICATING THRESHOLDS HIGHER THAN EXPECTED OR A SUSPENSION. THE PATIENT HAD A PRE-SYNCOPAL EPISODE. FURTHER EVALUATION WAS RECOMMENDED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE. ADDITIONAL INFORMATION WAS RECEIVED, THE NON-BOSTON SCIENTIFIC RV LEAD WAS SURGICALLY ABANDONED AND REPLACED, THERE WERE NO ISSUES FOUND RELATED TO THIS PACEMAKER. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20973064&pc=LWP | 2124215-2024-80515 | 2024-12-11 05:00 | 2024-12-18 05:00 | Malfunction | Failure to Capture; High Capture Threshold | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS AN ATTEMPTED IMPLANT. DURING A DEVICE REPLACEMENT PROCEDURE, THIS PACEMAKER EXHIBITED INTERMITTENT LOSS OF CAPTURE (LOC) WITH HIGH PACING THRESHOLDS. THE WAS SUCCESSFULLY COMPLETED WITH ANOTHER PACEMAKER. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. AT THIS TIME, THIS PACEMAKER HAS NOT BEEN RETURNED TO BOSTON SCIENTIFIC FOR FURTHER ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21145895&pc=LWP | 2124215-2025-01170 | 2024-12-20 05:00 | 2025-01-13 05:00 | Malfunction | Failure to Capture; High impedance | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED HIGH OUT OF RANGE PACING IMPEDANCE MEASUREMENTS ON THE RIGHT VENTRICULAR (RV) CHANNEL AND LOSS OF CAPTURE (LOC). THE PATIENT REMAINED UNDER MONITORING. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21214632&pc=LWP | 2124215-2025-03100 | 2025-01-02 05:00 | 2025-01-23 05:00 | Malfunction | Failure to Capture; High impedance; High Capture Threshold | ACCOLADE MRI SR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT DURING IMPLANT PROCEDURE THIS PACEMAKER SYSTEM EXHIBITED HIGH OUT OF RANGE PACING IMPEDANCE MEASUREMENTS AND HIGH PACING THRESHOLDS. TECHNICAL SERVICES (TS) SUSPECTED POSSIBLE LOSS OF CAPTURE (LOC) BUT IT WAS NOT CONCLUSIVE. THE PATIENT REMAINED UNDER MONITORING. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21283423&pc=LWP | 2124215-2025-05018 | 2025-01-16 05:00 | 2025-01-31 05:00 | Malfunction | Failure to Capture; Over-Sensing | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED LOSS OF CAPTURE (LOC) ON THE VENTRICULAR CHANNEL. ADDITIONALLY, FARFIELD OVERSENSING OF VENTRICULAR SIGNALS ON THE ATRIAL CHANNEL. ISOMETRICS WERE PERFORMED, AND NO NOISE WAS OBSERVED. ALL ELECTRICAL MEASUREMENTS WERE WITHIN NORMAL RANGES. TECHNICAL SERVICES (TS) SUGGESTED REPROGRAMMING. THE PHYSICIAN WILL CONTINUE TO MONITOR THE PATIENT. THE DEVICE REMAINS IN USE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21237285&pc=LWP | 2124215-2025-03647 | 2024-10-10 04:00 | 2025-01-26 05:00 | Malfunction | Failure to Capture; Over-Sensing | ACCOLADE SR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER HAD STORED SIGNAL ARTIFACT MONITORING (SAM) EPISODE AS WELL AS SUSPECTED LOSS OF CAPTURE (LOC) ON THE RIGHT VENTRICULAR (RV) LEAD CHANNEL. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE.Event Description: IT WAS REPORTED THAT THIS PACEMAKER HAD STORED SIGNAL ARTIFACT MONITORING (SAM) EPISODE AS WELL AS SUSPECTED LOSS OF CAPTURE (LOC) ON THE RIGHT VENTRICULAR (RV) LEAD CHANNEL. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21000069&pc=LWP | 2124215-2024-81400 | 2024-12-13 05:00 | 2024-12-23 05:00 | Malfunction | Failure to Capture; Over-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED ATRIAL CHANNEL OVERSENSING OF VENTRICULAR SIGNALS AS WELL AS POSSIBLE LOSS OF CAPTURE (LOC), AS THE ELECTROGRAM (EGM) HAD AN ODD APPEARANCE THAT SUGGESTED POTENTIAL LOC. IT WAS NOTED THAT THE PRESENTING EGMS HAVE ALWAYS APPEARED THIS WAY. THIS PACEMAKER REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21021154&pc=LWP | 2124215-2024-81984 | 2024-12-14 05:00 | 2024-12-26 05:00 | Malfunction | Failure to Capture; Over-Sensing; High Capture Threshold | ACCOLADE EL DR | Bradycardia | N970003 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS PACEMAKER SYSTEM PRESENTED TO THE EMERGENCY ROOM (ER). UPON EXAMINATION, THE PATIENT EXPERIENCED BRADYCARDIA WITH A PULSE IN THE 30'S RANGE. TECHNICAL SERVICES REVIEWED DEVICE DATA AND DISCUSSED THAT SOME OF THE VENTRICULAR PACING BEATS LOOK DIFFERENT AND COULD BE FUSION OR RIGHT VENTRICULAR (RV) LOSS OF CAPTURE. ADDITIONALLY, THRESHOLDS WERE HIGH MEASURING 2.7V@0.4MS. REVIEW OF THE PRESENTING ELECTROGRAM (EGM) SHOWS THERE ARE PREMATURE ATRIAL CONTRACTIONS (PAC) AND PREMATURE VENTRICULAR CONTRACTIONS. HOWEVER, SOME OF THE PACS APPEAR TO BE FAR-FIELD OVERSENSING THAT FALL IN REFRACTORY OR BLANKING. THE PATIENT IS BEING SENT TO THE HOSPITAL DUE TO OTHER HEALTH ISSUES WHERE HER ELECTROPHYSIOLOGY (EP) PHYSICIAN IS. AT THIS TIME, THE DEVICE AND NON-BOSTON SCIENTIFIC RV LEAD REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20953463&pc=LWP | 2124215-2024-79883 | 2024-11-19 05:00 | 2024-12-17 05:00 | Malfunction | Failure to Capture; Pacing Problem; Failure to Read Input Signal | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM STORED AN EPISODE OF PACEMAKER MEDIATED TACHYCARDIA (PMT) EPISODES. TECHNICAL SERVICES (TS) WAS CONSULTED AND DISCUSSED THE STORED EPISODES, WITH LOSS OF CAPTURE (LOC) SUSPECTED. ADDITIONALLY, LOW AMPLITUDE WAS NOTED FOR THE RIGHT VENTRICULAR (RV). THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21175299&pc=LWP | 2124215-2025-01982 | 2024-12-25 05:00 | 2025-01-16 05:00 | Malfunction | Failure to Capture; Pacing Problem; High Capture Threshold | ACCOLADE MRI EL DR | Bradycardia | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM STORED AN EPISODE OF PACEMAKER MEDIATED TACHYCARDIA (PMT). THE PMT SEEMED TO BE SINUS DRIVEN AND THE ALGORITHM SUCCESSFULLY TERMINATED THE EPISODE. IN ADDITION, LOSS OF CAPTURE (LOC) WAS SUSPECTED AND AN INCREASED THRESHOLD ON THE RIGHT VENTRICULAR (RV) LEAD WAS OBSERVED. THE PATIENT PRESENTED BRADYCARDIA. FURTHER MONITORING WAS RECOMMENDED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE. Manufacturer Narrative: THE FIRST NAME OF THE INITIAL REPORTER WAS NOT PROVIDED TO BSC. THE LOCAL AREA SALES REPRESENTATIVE WAS CONTACTED FOR ADDITIONAL INFORMATION. AT THIS TIME, NO FURTHER INFORMATION IS AVAILABLE. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THIS REPORT WILL BE UPDATED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21027921&pc=LWP | 2124215-2024-82163 | 2024-12-02 05:00 | 2024-12-27 05:00 | Malfunction | Failure to Capture; Under-Sensing | ACCOLADE MRI EL DR | Fatigue | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM WAS SUSPECTED OF EXHIBITING INTERMITTENT LOSS OF RIGHT ATRIAL (RA) LEAD CAPTURE. THE PATIENT REPORTED FATIGUE SINCE SYSTEM IMPLANT. TECHNICAL SERVICES (TS) WAS CONSULTED AND PROVIDED REPROGRAMMING OPTIONS AS CROSS CHAMBER BLANKING WAS NOTED. TS PROVIDED EVALUATION OPTIONS. THIS PACEMAKER SYSTEM REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20916354&pc=LWP | 2124215-2024-78780 | 2024-11-30 05:00 | 2024-12-12 05:00 | Malfunction | Failure to Capture; Under-Sensing | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE PACEMAKER EXHIBITED FUNCTIONAL UNDERSENSING AS PREMATURE VENTRICULAR CONTRACTIONS (PVCS) WERE FALLING INTO THE BLANKING PERIOD, AND LOSS OF CAPTURE AS THE DEVICE WOULD PACE AFTER THE PVCS BEFORE THE HEART REPOLARIZED. TECHNICAL SERVICES (TS) WAS CONTACTED, AND TS DISCUSSED PROGRAMMING OPTIONS. THE PACEMAKER SYSTEM REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21267783&pc=LWP | 2124215-2025-04539 | 2025-01-12 05:00 | 2025-01-29 05:00 | Malfunction | Failure to Read Input Signal; Under-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER ALERTED FOR INTRINSIC AMPLITUDE OUT OF RANGE ON THE RIGHT VENTRICULAR (RV) LEAD. THE R WAVE MEASURED AT 1.0MV (MILLIVOLTS) THIS WEEK. THE STORED ELECTROGRAM (EGM) SHOWED AN ISOLATED UNDERSENSED R WAVE AT AUTOMATIC GAIN CONTROL (AGC) 0.6MV. FURTHER REVIEW OF PROGRAMMED PARAMETERS WAS RECOMMENDED. THIS ALERT WILL NOT RETRIGGER UNTIL THE DEVICE IS INTERROGATED WITH A PROGRAMMER. THE DEVICE AND THE NON-BOSTON SCIENTIFIC RV LEAD REMAIN IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21212269&pc=LWP | 2124215-2025-03034 | 2025-01-19 05:00 | 2025-01-22 05:00 | Malfunction | Failure to Read Input Signal; Under-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS DEVICE GENERATED THREE ALERTS FOR VENTRICULAR TACHYCARDIA (VT) EPISODES. ELECTROGRAMS (EGM) FROM THE MOST RECENT EVENT SHOWED AN ATRIAL ARRHYTHMIA WITH 1:1 CONDUCTION AND AN AVERAGE VENTRICULAR RATE OF 161-176BPM. ATRIAL SENSITIVITY IS PROGRAMMED TO FIXED OUTPUT AT 0.5MV AND REVIEW OF THE STORED DATA NOTED ATRIAL INTRINSIC AMPLITUDES VARIED FROM 0.3MV TO 6.7MV. ADDITIONALLY, FREQUENT ATRIAL UNDERSENSING WAS OBSERVED IN RYTHMIQ EPISODES. THE CLINICAL OBSERVATIONS WERE DOCUMENTED AND THE DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: AS NO FURTHER INFORMATION CONCERNING THIS REPORT IS EXPECTED, OUR INVESTIGATION IS COMPLETE. THIS INVESTIGATION WILL BE UPDATED SHOULD FURTHER INFORMATION BE PROVIDED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21204248&pc=LWP | 2124215-2025-02794 | 2024-12-26 05:00 | 2025-01-21 05:00 | Malfunction | Failure to Read Input Signal; Under-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER INTRINSIC AMPLITUDE IS OUT OF RANGE ON THE RIGHT ATRIAL (RA) LEAD. THE PRESENTING ELECTROGRAM (EGM) SHOWED ISOLATED ATRIAL UNDERSENSING. IT WAS RECOMMENDED TO REVIEW THE PROGRAMMED ATRIAL SENSITIVITY TO ENSURE SENSING AND MAINTENANCE OF ATRIOVENTRICULAR (AV) SYNCHRONY. THIS ALERT WILL NOT RETRIGGER UNTIL THE DEVICE IS INTERROGATED WITH A PROGRAMMER. AT THIS TIME, THE DEVICE AND LEAD REMAIN IN-SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21123760&pc=LWP | 2124215-2024-83609 | 2024-12-15 05:00 | 2025-01-10 05:00 | Malfunction | Failure to Read Input Signal; Under-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THE PACEMAKER ALERTED DUE TO ATRIAL ARRHYTHMIA BURDEN. LOW ATRIAL INTRINSIC AMPLITUDE WITH A SINGLE UNDERSENSED ATRIAL SIGNAL ON THE PRESENTING ELECTROGRAM WAS ALSO OBSERVED. THE ATRIAL SENSITIVITY WAS PROGRAMMED TO AUTOMATIC GAIN CONTROL AT 0.25 MV. TECHNICAL SERVICES RECOMMENDED REVIEW OF THE PROGRAMMED PARAMETERS. THE PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21129666&pc=LWP | 2124215-2025-00788 | 2022-05-11 04:00 | 2025-01-10 05:00 | Malfunction | Failure to Read Input Signal; Under-Sensing | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED INTERMITTENT UNDER SENSING ON THE RIGHT VENTRICULAR (RV) LEAD CHANNEL. TECHNICAL SERVICES (TS) WAS CONSULTED AND A DECREASE IN AMPLITUDE WAS NOTED, ADJUSTING THE SENSITIVITY OF THE RV CHANNEL WAS RECOMMENDED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20984578&pc=LWP | 2124215-2024-80882 | 2024-11-23 05:00 | 2024-12-19 05:00 | Malfunction | Failure to Read Input Signal; Under-Sensing | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER INTRINSIC AMPLITUDE IS OUT OF RANGE AND A VENTRICULAR TACHYCARDIA (VT) EPISODE WAS RECORDED. THE PRESENTING ELECTROGRAM (EGM) SHOW ATRIAL ARRHYTHMIA WITH RAPID VENTRICULAR RESPONSE (RVR) WITH SIGNIFICANT ATRIAL UNDERSENSING RESULTING IN INCORRECT DETECTION OF VT. THE ATRIAL INTRINSIC AMPLITUDE TREND SHOWS A RANGE BETWEEN 0.3MV TO 1.8MV (MILLIVOLTS). THE ATRIAL SENSITIVITY IS PROGRAMMED TO AUTOMATIC GAIN CONTROL (AGC) 0.4MV. FURTHER REVIEW OF ATRIAL SENSITIVITY PROGRAMMING WAS RECOMMENDED TO ENSURE APPROPRIATE RHYTHM DETECTION. THE DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20947517&pc=LWP | 2124215-2024-79599 | 2024-11-25 05:00 | 2024-12-16 05:00 | Malfunction | Failure to Read Input Signal; Under-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER RECORDED AN ALERT FOR RIGHT VENTRICULAR (RV) AUTOMATIC THRESHOLD DETECTED AS SUSPENDED DUE TO LOW EVOKED RESPONSE. IN ADDITION, ATRIAL TACHY RESPONSE (ATR) EPISODES SHOW EVIDENCE OF ATRIAL UNDERSENSING. IN-OFFICE ASSESSMENT WAS RECOMMENDED. THE DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21213955&pc=LWP | 2124215-2025-03053 | 2024-09-02 04:00 | 2025-01-22 05:00 | Malfunction | False Positive Result; Over-Sensing; Data Problem | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT A SIGNAL ARTIFACT MONITOR (SAM) EVENT RESULTED IN MINUTE VENTILATION (MV) SENSOR SWITCH AND MV BEING DISABLED. REVIEW OF THE STORED ELECTROGRAMS (EGMS) SHOWED ATRIAL FIBRILLATION (AF) WITH NOISE MARKERS. THE CLINICAL OBSERVATIONS WERE DOCUMENTED. THE DEVICE REMAINS IN SERVICE AND NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS DEVICE HAS NOT BEEN RETURNED; THEREFORE, TECHNICAL ANALYSIS CANNOT BE CONDUCTED. INVESTIGATION OF THE AVAILABLE INFORMATION DETERMINED THIS DEVICE EXHIBITED INCORRECT CATEGORIZATION OF ATRIAL ARRHYTHMIA(S) AS MV NOISE. PLEASE SEE THE EVENT DESCRIPTION FIELD FOR MORE INFORMATION REGARDING THE SPECIFIC CIRCUMSTANCES OF THIS EVENT. A CORRECTIVE AND PREVENTIVE ACTION (CAPA) INVESTIGATION WAS PERFORMED TO FURTHER EVALUATE THIS BEHAVIOR AND IT WAS DETERMINED THAT A CORRECTIVE ACTION IS NOT WARRANTED AT THIS TIME. BOSTON SCIENTIFIC WILL CONTINUE TO MONITOR AND TREND THESE COMPLAINTS TO ENSURE THAT THE RATE REMAINS WITHIN EXPECTATIONS AS OUTLINED IN OUR QUALITY SYSTEMS PROCESS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20933545&pc=LWP | 2124215-2024-79206 | 2024-12-02 05:00 | 2024-12-14 05:00 | Malfunction | False Positive Result; Pacing Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER RECENTLY EXHIBITED ATRIAL OVER-SENSING, WHICH RESULTED IN PACING INHIBITION AND A DECREASED PACING RATE IN THIS PACEMAKER-DEPENDENT PATIENT. THE PHYSICIAN REPROGRAMMED THE DEVICE SENSITIVITY AND PERFORMED EXERCISE TESTING, WHICH REVEALED APPROPRIATE ATRIAL SENSING. AT THIS TIME, THE PACEMAKER REMAINS IMPLANTED AND IN-SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER RECENTLY EXHIBITED ATRIAL OVER-SENSING, WHICH RESULTED IN PACING INHIBITION AND A DECREASED PACING RATE IN THIS PACEMAKER-DEPENDENT PATIENT. THE PHYSICIAN REPROGRAMMED THE DEVICE SENSITIVITY AND PERFORMED EXERCISE TESTING, WHICH REVEALED APPROPRIATE ATRIAL SENSING. AT THIS TIME, THE PACEMAKER REMAINS IMPLANTED AND IN-SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS REPORT IS BEING SENT TO PROVIDE NEW INFORMATION IN SECTIONS H6 (EVALUATION CONCLUSION CODE) AND H11 (ADDITIONAL MFR. NARRATIVE). THIS DEVICE HAS NOT BEEN RETURNED; THEREFORE, TECHNICAL ANALYSIS CANNOT BE CONDUCTED. INVESTIGATION OF THE AVAILABLE INFORMATION DETERMINED THIS DEVICE EXHIBITED INCORRECT CATEGORIZATION OF ATRIAL ARRHYTHMIA(S) AS MV NOISE. PLEASE SEE THE B5 FIELD FOR MORE INFORMATION REGARDING THE SPECIFIC CIRCUMSTANCES OF THIS EVENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21270469&pc=LWP | 2124215-2025-04676 | 2025-01-15 05:00 | 2025-01-29 05:00 | Malfunction | High impedance | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT AT A FOLLOW-UP APPOINTMENT IN 2023, THE PHYSICIAN OBSERVED A SINGLE ELEVATED, OUT-OF-RANGE PACING IMPEDANCE MEASUREMENT (GREATER THAN 3,000 OHMS) FROM THE RIGHT VENTRICULAR (RV) LEAD. AT THAT TIME, THE PHYSICIAN PERFORMED REPROGRAMMING AND ENROLLED THE PATIENT IN REMOTE MONITORING. AS THE PATIENT IS NEARING A PROPHYLACTIC PACEMAKER EXPLANT PROCEDURE, BOSTON SCIENTIFIC TECHNICAL SERVICES (TS) WAS CONTACTED FOR A DATA REVIEW TO CONFIRM THE INTEGRITY OF THE RV LEAD. DURING THE REVIEW, TS NOTED ONLY ONE OTHER SIMILAR INSTANCE OF OUT-OF-RANGE IMPEDANCE (GREATER THAN 3,000 OHMS) IN (B)(6) 2024, AND SINCE THE PROGRAMMING TO UNIPOLAR, THE PACING IMPEDANCE MEASUREMENTS WERE STABLE AND IN-RANGE (BETWEEN 350 AND 450 OHMS). NEVERTHELESS, TS RECOMMENDED PERFORMING AN IN-CLINIC ASSESSMENT TO VERIFY THE LEAD INTEGRITY VIA PROVOCATIVE MANEUVERS, ISOMETRICS, AND DIAGNOSTIC IMAGING. SHOULD ANY ABNORMALITIES BE OBSERVED, IT WAS RECOMMENDED TO REPLACE THE LEAD DURING THE UPCOMING PROPHYLACTIC EXPLANT PROCEDURE. AT THIS TIME, THE PACEMAKER REMAINS IMPLANTED AND IN-SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: AT THIS TIME, NO FURTHER INFORMATION IS AVAILABLE. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THIS REPORT WILL BE UPDATED. IF WE LATER RECEIVE THE INFORMATION, WE WILL PROVIDE A FOLLOW UP REPORT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21257232&pc=LWP | 2124215-2025-04213 | 2025-01-13 05:00 | 2025-01-28 05:00 | Malfunction | High impedance | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED PACING IMPEDANCE HIGH OUT OF RANGE. PATIENT IS PACEMAKER DEPENDENT. THE DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21237284&pc=LWP | 2124215-2025-03645 | 2024-07-19 04:00 | 2025-01-26 05:00 | Malfunction | High impedance | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM EXHIBITED HIGH OUT OF RANGE PACING IMPEDANCES MEASURING ABOVE 3000 OHMS ON THE RIGHT VENTRICULAR (RV) LEAD CHANNEL, WHICH LED TO A LEAD SAFETY SWITCH (LSS). NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS SYSTEM REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21232539&pc=LWP | 2124215-2025-03508 | 2025-01-08 05:00 | 2025-01-24 05:00 | Malfunction | High impedance | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED HIGH OUT OF RANGE PACE IMPEDANCE MEASUREMENTS ON THE RIGHT VENTRICULAR (RV) LEAD. TECHNICAL SERVICES (TS) ADVISED THE HEALTH CARE PROFESSIONAL (HCP) TO FURTHER EVALUATE THE PATIENT. NO ADVERSE PATIENT EFFECTS WERE REPORTED. AT THIS TIME, THIS DEVICE SYSTEM REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21232540&pc=LWP | 2124215-2025-03509 | 2024-12-26 05:00 | 2025-01-24 05:00 | Malfunction | High impedance | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM TRIGGERED A LEAD SAFETY SWITCH (LSS), SWITCHING FROM BIPOLAR TO UNIPOLAR MODE DUE TO HIGH OUT-OF-RANGE RIGHT ATRIAL (RA) LEAD PACING IMPEDANCES EXCEEDING 2000 OHMS. THE IMPEDANCES MEASUREMENTS WERE REPORTED TO BE VARIABLE FROM 400 OHMS TO HIGHER THAN 2000 OHMS, AND NO NOISE WAS FOUND DURING TESTING. TECHNICAL SERVICES (TS) REVIEWED THE EVENT AND INDICATED THAT A SPRING CONTACT PROBLEM WAS SUSPECTED. THE RA LEAD IS A NON-BOSTON SCIENTIFIC DEVICE. TS ALSO PROVIDED REPROGRAMMING OPTIONS. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21203426&pc=LWP | 2124215-2025-02753 | 2024-12-31 05:00 | 2025-01-21 05:00 | Malfunction | High impedance | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED ABRUPT HIGH OUT OF RANGE PACE IMPEDANCE MEASUREMENTS ON THE RIGHT ATRIAL (RA) LEAD RESULTING IN A LEAD SAFETY SWITCH. THE INVOLVED LEAD IS A NON BOSTON SCIENTIFIC PRODUCT. A DEVICE CHANGEOUT PROCEDURE WAS PERFORMED AND THE DEVICE WAS EXPLANTED AND REPLACED PREVENTIVELY. THE RA LEAD REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE HAS BEEN RECEIVED FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21161802&pc=LWP | 2124215-2025-01538 | 2024-07-02 04:00 | 2025-01-15 05:00 | Malfunction | High impedance | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER RECORDED A RIGHT ATRIAL (RA) LEAD SAFETY SWITCH DUE TO OUT OF RANGE IMPEDANCE MEASUREMENTS GREATER THAN 3000 OHMS. THE LEAD WHEN PROGRAMMED TO UNIPOLAR THE IMPEDANCE MEASUREMENTS WAS WITH IN RANGE. THE CLINIC WILL CONTINUE TO MONITOR THE PROGRESS, AND A LEAD REVISION WILL BE CONSIDERED IF THERE IS AN ADDITIONAL SIGNIFICANT IMPEDANCE CHANGE. THIS PACEMAKER REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21138909&pc=LWP | 2124215-2025-01003 | 2024-11-10 05:00 | 2025-01-13 05:00 | Malfunction | High impedance | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM TRIGGERED A LEAD SAFETY SWITCH (LSS) DUE TO A HIGH OUT OF RANGE RIGHT VENTRICULAR (RV) LEAD PACING IMPEDANCE MEASUREMENT OF GREATER THAN 2000 OHMS. THE HEALTH CARE PROFESSIONAL HAD CALLED IN FOR ASSISTANCE WITH HAVING THIS PACEMAKER DEVICE INTO MAGNETIC RESONANCE IMAGING (MRI) MODE. TECHNICAL SERVICES (TS) CONFIRMED THAT THIS SYSTEM FOR MRI COMPATIBLE AND ALL LEAD MEASUREMENTS WERE CURRENTLY WITHIN NORMAL RANGE. THE BOSTON SCIENTIFIC REPRESENTATIVE HAD CALLED BACK CONFIRMING THE EXIT OF MRI MODE AND STATED THAT MOST RECENT LEAD IMPEDANCE TESTING WAS WITHIN NORMAL LIMITS. THE DEVICE APPEARED TO BE FUNCTIONING AS PROGRAMMED AND REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21106883&pc=LWP | 2124215-2025-00313 | 2024-12-17 05:00 | 2025-01-08 05:00 | Malfunction | High impedance | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED RIGHT ATRIAL (RA) LEAD AND RIGHT VENTRICULAR (RV) LEAD IMPEDANCE HIGH OUT OF RANGE. THE PATIENT HAD UNDERGONE A ATRIAL FIBRILLATION (AF) ABLATION AT THE TYPE OF THE ISSUE. THE MEASUREMENTS CAME BACK TO NORMAL RANGE AFTER THE PROCEDURE. TECHNICAL SERVICES (TS) WAS CONSULTED AND EXPLAINED THERE WAS NO NEED TO REPROGRAM. THE DEVICE REMAINS IN SERVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21056429&pc=LWP | 2124215-2025-00029 | 2024-12-23 05:00 | 2025-01-02 05:00 | Malfunction | High impedance | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE HEALTH CARE PROFESSIONAL (HCP) HAD CALLED IN FOR PROGRAMMING THIS PACEMAKER SYSTEM IN MAGNETIC RESONANCE IMAGING (MRI) MODE. TECHNICAL SERVICES (TS) REVIEWED AND STATED THAT THIS WAS NOT MRI COMPATIBLE DUE TO RIGHT ATRIAL (RA) LEAD OUT OF RANGE PACING IMPEDANCE MEASUREMENTS, MEASURING GREATER THAN 3000 OHMS. TS RECOMMENDED TO CONSULT THE CARDIOLOGIST TO SEE HOW THEY WANTED TO PROCEED. THIS PACEMAKER DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21020462&pc=LWP | 2124215-2024-81965 | 2024-12-16 05:00 | 2024-12-26 05:00 | Malfunction | High impedance | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT ONE DAY AFTER AN ABLATION PROCEDURE, THIS PACEMAKER EXHIBITED HIGH OUT OF RANGE IMPEDANCE MEASUREMENTS ON THE RIGHT VENTRICULAR (RV) LEAD RESULTING IN A LEAD SAFETY SWITCH. THE HEALTH CARE PROFESSIONAL (HCP) STATED THIS PATIENT IS NOT SCHEDULED TO BE SEEN AGAIN UNTIL MARCH. TECHNICAL SERVICES (TS) RECOMMENDED PERFORMING FURTHER EVALUATION EARLIER IF THE PATIENT REPORTS SYMPTOMS. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20984581&pc=LWP | 2124215-2024-80893 | 2024-10-01 04:00 | 2024-12-19 05:00 | Malfunction | High impedance | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT WHILE PREPARING THIS PACEMAKER SYSTEM FOR A MAGNETIC RESONANCE IMAGING (MRI) SCAN, A DEVICE CHECK REVEALED THERE WAS A LEAD SAFETY SWITCH (LSS) DUE TO HIGH OUT-OF-RANGE PACE IMPEDANCE MEASUREMENTS GREATER THAN 3,000 OHMS ON THE RIGHT ATRIAL (RA) LEAD A COUPLE MONTHS PRIOR. THE HEALTH CARE PROFESSIONAL (HCP) WILL CONTACT CARDIOLOGY FOR AN UPDATED MRI ORDER. THIS PACEMAKER REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THE LOCAL AREA FIELD REPRESENTATIVE WAS CONTACTED FOR ADDITIONAL INFORMATION REGARDING EVENT CAUSE AND RESOLUTION. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20973063&pc=LWP | 2124215-2024-80518 | 2024-10-14 04:00 | 2024-12-18 05:00 | Malfunction | High impedance | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER DEVICE EXHIBITED A GRADUAL INCREASE IN PACING IMPEDANCE. THE IMPEDANCE WAS 2,023 OHMS IN (B)(6) 2024 AND ROSE TO 3,000 OHMS IN NOVEMBER. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. A TECHNICAL SERVICE (TS) CONSULTANT REVIEWED DEVICE DATA ADVISING OF LEAD SAFETY SWITCH FROM BIPOLAR TO UNIPOLAR, WITH NO SENSING ISSUES AND NO LOSS OF CAPTURE. TS ADVISED THAT THE RISE IN PACING IMPEDANCE COULD BE RELATED TO PHYSIOLOGIC REASON. THE PHYSICIAN WILL MONITOR THE PATIENT CLOSELY. THE DEVICE REMAINS IMPLANTED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20925731&pc=LWP | 2124215-2024-78989 | 2024-12-04 05:00 | 2024-12-13 05:00 | Malfunction | High impedance | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXHIBITING RIGHT VENTRICULAR (RV) LEAD PACING IMPEDANCE HIGH OUT OF RANGE. AN MAGNETIC RESONANCE IMAGING (MRI) WAS INDENTED TO BE PERFORMED; HOWEVER, MRI PROTECTION MODE WAS NOT AVAILABLE DUE TO RV SET TO UNIPOLAR/BIPOLAR. THE DEVICE REMAINS IN SERVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21020789&pc=LWP | 2124215-2024-81975 | 2024-12-02 05:00 | 2024-12-26 05:00 | Malfunction | High impedance; Off-Label Use | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT ASSISTANCE WAS REQUESTED TO PROGRAM THIS PACEMAKER INTO MAGNETIC RESONANCE IMAGING (MRI) MODE. TECHNICAL SERVICES (TS) CONFIRMED THAT THE SYSTEM IS MRI CONDITIONAL; HOWEVER, THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED HIGH OUT-OF-RANGE PACING IMPEDANCES, MEASURING ABOVE 2000 OHMS, WHICH TRIGGERED A LEAD SAFETY SWITCH (LSS). ADDITIONALLY, NOISE WAS OBSERVED IN THE ELECTROCARDIOGRAM (EGM). AS A RESULT, THE SYSTEM NO LONGER QUALIFIED AS MRI CONDITIONAL. SUBSEQUENTLY, THE PACEMAKER WAS PROGRAMMED INTO MRI PROTECTION MODE. UPON EXITING MRI MODE, IT WAS CONFIRMED THAT THE DEVICE WAS FUNCTIONING AS PROGRAMMED. NO ADVERSE PATIENT EFFECTS WERE REPORTED IN CONNECTION WITH THIS OFF-LABEL USE. THIS PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20951748&pc=LWP | 2124215-2024-79747 | 2024-12-04 05:00 | 2024-12-17 05:00 | Malfunction | High impedance; Over-Sensing | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXHIBITING PACING IMPEDANCE HIGH OUT OF RANGE. TECHNICAL SERVICES (TS) WAS CONSULTED AND NOTED MINUTE VENTILATION (MV) CHOP AND OVERSENSING. TS SUSPECTS A LEAD ISSUE AND RECOMMENDS FURTHER EVALUATION. THE DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXHIBITING PACING IMPEDANCE HIGH OUT OF RANGE. TECHNICAL SERVICES (TS) WAS CONSULTED AND NOTED MINUTE VENTILATION (MV) CHOP AND OVERSENSING. TS SUSPECTS A LEAD ISSUE AND RECOMMENDS FURTHER EVALUATION. THE DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS PRODUCT HAS NOT BEEN RETURNED TO BOSTON SCIENTIFIC, AND AS A RESULT, LABORATORY ANALYSIS COULD NOT BE CONDUCTED. INVESTIGATION OF THE AVAILABLE INFORMATION DETERMINED THIS DEVICE EXHIBITED OVERSENSING OF NOISE GENERATED BY THE MINUTE VENTILATION (MV)/RESPIRATORY SENSOR THAT IS RELATED TO INTERMITTENT INCREASES IN IMPEDANCE MEASUREMENTS. ADDITIONALLY, ENGINEERING ANALYSIS AND TESTING OF RETURNED PRODUCTS HAS IDENTIFIED THAT ANY REPEATED, SMALL MOVEMENTS OF THE LEAD TERMINAL RING CAN CREATE WEARING OF THE LEAD TERMINAL RING AND GENERATE MICROSCOPIC PARTICLES, WHICH MAY ACCUMULATE AND OXIDIZE OVER TIME. THIS CAN IMPACT THE CONNECTION BETWEEN THE SPRING CONTACT AND THE LEAD RING, RESULTING IN INTERMITTENT CHANGES IN IMPEDANCE MEASUREMENTS. PLEASE REFER TO THE DESCRIPTION FIELD FOR MORE INFORMATION REGARDING THE SPECIFIC CIRCUMSTANCES OF THIS EVENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20957157&pc=LWP | 2124215-2024-80032 | 2024-12-06 05:00 | 2024-12-17 05:00 | Malfunction | High impedance; Over-Sensing | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER RECORDED A LEAD SAFETY SWITCH DUE TO RIGHT ATRIAL (RA) LEAD PACE IMPEDANCE MEASUREMENTS GREATER THAN 3000 OHMS AND RA OVERSENSING ON PRESENTING ELECTROGRAM (EGM) WAS ALSO NOTED. TECHNICAL SERVICES (TS) DISCUSSED REPROGRAMMING OPTIONS AND RECOMMENDED FURTHER REVIEW WITH A BSC PROGRAMMER. THIS PACEMAKER REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THIS PACEMAKER RECORDED A LEAD SAFETY SWITCH DUE TO RIGHT ATRIAL (RA) LEAD PACE IMPEDANCE MEASUREMENTS GREATER THAN 3000 OHMS AND RA OVERSENSING ON PRESENTING ELECTROGRAM (EGM) WAS ALSO NOTED. TECHNICAL SERVICES (TS) DISCUSSED REPROGRAMMING OPTIONS AND RECOMMENDED FURTHER REVIEW WITH A BSC PROGRAMMER. THIS PACEMAKER REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS PRODUCT HAS NOT BEEN RETURNED TO BOSTON SCIENTIFIC, AND AS A RESULT, LABORATORY ANALYSIS COULD NOT BE CONDUCTED. THE ASSOCIATED INVESTIGATION DETERMINED THAT THIS DEVICE EXHIBITED INTERMITTENT CHANGES IN IMPEDANCE MEASUREMENTS WITH NO CONCLUSIVE EVIDENCE OF A PRODUCT PERFORMANCE ISSUE OR INADEQUATE LEAD-TO-DEVICE CONNECTION; PLEASE REFER TO THE DESCRIPTION FOR MORE INFORMATION REGARDING THE SPECIFIC CIRCUMSTANCES OF THIS EVENT. ENGINEERING ANALYSIS AND TESTING OF RETURNED PRODUCTS HAS DETERMINED THAT REPEATED, SMALL MOVEMENTS OF THE LEAD TERMINAL RING ALONG WITH VARIANCES IN THE CONNECTION CONTACT FORCE AND SURFACE ROUGHNESS CAN IMPACT THE CONNECTION BETWEEN THE SPRING CONTACT AND THE LEAD RING, RESULTING IN INTERMITTENT CHANGES IN IMPEDANCE MEASUREMENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21044949&pc=LWP | 2124215-2024-82511 | 2024-11-11 05:00 | 2024-12-30 05:00 | Malfunction | High impedance; Over-Sensing; Data Problem | ACCOLADE MRI SR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT ALERTS WERE GENERATED FOR A RIGHT VENTRICULAR (RV) PACING IMPEDANCE MEASUREMENT GREATER THAN 3000 OHMS AND A SIGNAL ARTIFACT MONITOR (SAM) DEVICE DIAGNOSTIC NOT ENABLED. THE OUT-OF-RANGE PACING IMPEDANCE CAUSED THE LEAD SAFETY SWITCH (LSS) TO CHANGE LEAD CONFIGURATION FROM BIPOLAR TO UNIPOLAR CONFIGURATION. ADDITIONALLY, THE MINUTE VENTILATION (MV) SENSOR WAS DISABLED BY THE SIGNAL ARTIFACT MONITOR (SAM) DEVICE DIAGNOSTIC DUE TO MV ARTIFACT. THIS SYSTEM CONSISTS OF A BOSTON SCIENTIFIC IMPLANTABLE DEVICE AND A COMPETITIVE LEAD. THE CLINICAL OBSERVATIONS WERE DOCUMENTED, AND THE SYSTEM REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT ALERTS WERE GENERATED FOR A RIGHT VENTRICULAR (RV) PACING IMPEDANCE MEASUREMENT GREATER THAN 3000 OHMS AND A SIGNAL ARTIFACT MONITOR (SAM) DEVICE DIAGNOSTIC NOT ENABLED. THE OUT-OF-RANGE PACING IMPEDANCE CAUSED THE LEAD SAFETY SWITCH (LSS) TO CHANGE LEAD CONFIGURATION FROM BIPOLAR TO UNIPOLAR CONFIGURATION. ADDITIONALLY, THE MINUTE VENTILATION (MV) SENSOR WAS DISABLED BY THE SIGNAL ARTIFACT MONITOR (SAM) DEVICE DIAGNOSTIC DUE TO MV ARTIFACT. THIS SYSTEM CONSISTS OF A BOSTON SCIENTIFIC IMPLANTABLE DEVICE AND A COMPETITIVE LEAD. THE CLINICAL OBSERVATIONS WERE DOCUMENTED, AND THE SYSTEM REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: AS NO FURTHER INFORMATION CONCERNING THIS REPORT IS EXPECTED, OUR INVESTIGATION IS COMPLETE. THIS INVESTIGATION WILL BE UPDATED SHOULD FURTHER INFORMATION BE PROVIDED. Manufacturer Narrative: THIS PRODUCT HAS NOT BEEN RETURNED TO BOSTON SCIENTIFIC, AND AS A RESULT, LABORATORY ANALYSIS COULD NOT BE CONDUCTED. THE ASSOCIATED INVESTIGATION DETERMINED THAT THIS DEVICE EXHIBITED INTERMITTENT CHANGES IN IMPEDANCE MEASUREMENTS WITH NO CONCLUSIVE EVIDENCE OF A PRODUCT PERFORMANCE ISSUE OR INADEQUATE LEAD-TO-DEVICE CONNECTION; PLEASE REFER TO THE EVENT DESCRIPTION FOR MORE INFORMATION REGARDING THE SPECIFIC CIRCUMSTANCES OF THIS EVENT. ENGINEERING ANALYSIS AND TESTING OF RETURNED PRODUCTS HAS DETERMINED THAT REPEATED, SMALL MOVEMENTS OF THE LEAD TERMINAL RING ALONG WITH VARIANCES IN THE CONNECTION CONTACT FORCE AND SURFACE ROUGHNESS CAN IMPACT THE CONNECTION BETWEEN THE SPRING CONTACT AND THE LEAD RING, RESULTING IN INTERMITTENT CHANGES IN IMPEDANCE MEASUREMENT. AS NO FURTHER INFORMATION CONCERNING THIS REPORT IS EXPECTED, OUR INVESTIGATION IS COMPLETE. THIS INVESTIGATION WILL BE UPDATED SHOULD FURTHER INFORMATION BE PROVIDED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21270472&pc=LWP | 2124215-2025-04621 | 2025-01-05 05:00 | 2025-01-29 05:00 | Malfunction | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER REVERTED TO SAFETY MODE. REPLACEMENT WAS HIGHLY RECOMMENDED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE. ADDITIONAL INFORMATION WAS PROVIDED, THIS PACEMAKER WAS EXPLANTED AND REPLACED. OTHER THAN THE PROCEDURE, NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. AT THIS TIME, THIS PACEMAKER HAS NOT BEEN RETURNED TO BOSTON SCIENTIFIC FOR FURTHER ANALYSIS.Event Description: IT WAS REPORTED THAT THIS PACEMAKER REVERTED TO SAFETY MODE. REPLACEMENT WAS HIGHLY RECOMMENDED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21242840&pc=LWP | 2124215-2025-03794 | 2025-01-08 05:00 | 2025-01-27 05:00 | Malfunction | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS FOUND TO BE IN SAFETY MODE. TECHNICAL SERVICES (TS) RECOMMENDED URGENT REPLACEMENT. THIS PACEMAKER REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21202990&pc=LWP | 2124215-2025-02748 | 2025-01-07 05:00 | 2025-01-21 05:00 | Malfunction | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS FOUND TO BE OPERATING IN SAFETY MODE DURING A ROUTINE FOLLOW-UP. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21175290&pc=LWP | 2124215-2025-01958 | 2024-12-24 05:00 | 2025-01-16 05:00 | Malfunction | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER REVERTED TO SAFETY MODE. IT WAS NOTED THAT THE DEVICE RECENTLY SHOWED A BATTERY STATUS OF 1.5 YEARS REMAINING. TECHNICAL SERVICES (TS) RECOMMENDED DEVICE REPLACEMENT. THE SPECIFIC DEVICE SERIAL NUMBER, DATE OF IMPLANT AND FACILITY OR PHYSICIAN INFORMATION ARE UNKNOWN AT THIS TIME, HOWEVER, HAVE BEEN REQUESTED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. AVAILABLE INFORMATION INDICATES THIS DEVICE REMAINS IN SERVICE. Manufacturer Narrative: THE LOCAL AREA SALES REPRESENTATIVE WAS CONTACTED FOR ADDITIONAL INFORMATION. BECAUSE THE PRODUCT SERIAL NUMBER IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) # AND OTHER SPECIFIC PRODUCT INFORMATION. AT THIS TIME, NO FURTHER INFORMATION IS AVAILABLE. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE THIS REPORT WILL BE UPDATED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21161804&pc=LWP | 2124215-2025-01528 | 2025-01-07 05:00 | 2025-01-15 05:00 | Malfunction | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI SR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS FOUND TO HAVE REVERTED TO A SAFETY MODE STATUS. RHYTHM CARE THEN RECOMMENDED DEVICE REPLACEMENT. THIS DEVICE CURRENTLY REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21152268&pc=LWP | 2124215-2025-01334 | 2024-12-28 05:00 | 2025-01-14 05:00 | Malfunction | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS FOUND TO BE IN SAFETY MODE THREE WEEKS AFTER IT WAS MANUFACTURED. NO PATIENT INVOLVEMENT WAS REPORTED SINCE IT WAS NEVER IN SERVICE. THIS DEVICE HAS BEEN RECEIVED FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21027913&pc=LWP | 2124215-2024-82157 | 2024-12-18 05:00 | 2024-12-27 05:00 | Malfunction | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS FOUND TO BE IN SAFETY MODE. TECHNICAL SERVICES (TS) WAS CONSULTED AND RECOMMENDED DEVICE REPLACEMENT. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20973747&pc=LWP | 2124215-2024-80507 | 2024-11-25 05:00 | 2024-12-19 05:00 | Malfunction | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT DURING A RETROSPECTIVE REVIEW OF DEVICE DATA IT WAS IDENTIFIED THAT THIS PACEMAKER WAS FOUND TO BE IN SAFETY MODE. NO ADDITIONAL DETAILS WERE PROVIDED. AT THIS POINT, THE DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20984150&pc=LWP | 2124215-2024-80878 | 2024-12-13 05:00 | 2024-12-19 05:00 | Malfunction | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE PACEMAKER | No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: IT WAS REPORTED THAT DURING A DISCUSSION WITH THE PHYSICIAN, IT WAS MENTIONED THAT UP TO THREE OF THEIR PATIENTS' IMPLANTABLE ACCOLADE PACEMAKERS HAD ENTERED SAFETY MODE. THE SPECIFIC DEVICE DETAILS ARE CURRENTLY UNKNOWN. AT THIS TIME, THIS PACEMAKER REMAINS IMPLANTED AND IN-SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20984580&pc=LWP | 2124215-2024-80879 | 2024-12-13 05:00 | 2024-12-19 05:00 | Malfunction | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE PACEMAKER | No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: IT WAS REPORTED THAT DURING A DISCUSSION WITH THE PHYSICIAN, IT WAS MENTIONED THAT UP TO THREE OF THEIR PATIENTS' IMPLANTABLE ACCOLADE PACEMAKERS HAD ENTERED SAFETY MODE. THE SPECIFIC DEVICE DETAILS ARE CURRENTLY UNKNOWN. AT THIS TIME, THIS PACEMAKER REMAINS IMPLANTED AND IN-SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20984584&pc=LWP | 2124215-2024-80883 | 2024-12-13 05:00 | 2024-12-19 05:00 | Malfunction | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE PACEMAKER | No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: IT WAS REPORTED THAT DURING A DISCUSSION WITH THE PHYSICIAN, IT WAS MENTIONED THAT UP TO THREE OF THEIR PATIENTS' IMPLANTABLE ACCOLADE PACEMAKERS HAD ENTERED SAFETY MODE. THE SPECIFIC DEVICE DETAILS ARE CURRENTLY UNKNOWN. AT THIS TIME, THIS PACEMAKER REMAINS IMPLANTED AND IN-SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20968394&pc=LWP | 2124215-2024-80362 | 2024-12-16 05:00 | 2024-12-18 05:00 | Malfunction | Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER UPON INTERROGATION WAS FOUND TO BE IN SAFETY MODE. DEVICE REPLACEMENT WAS RECOMMENDED. AT THIS TIME, THE PACEMAKER REMAINS IN-SERVICE AND THERE IS NO EVIDENCE TO SUGGEST A DEVICE REPLACEMENT PROCEDURE HAS BEEN SCHEDULED AT THIS TIME. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21048180&pc=LWP | 2124215-2024-82637 | 2024-12-08 05:00 | 2024-12-31 05:00 | Malfunction | Incorrect, Inadequate or Imprecise Result or Readings; Telemetry Discrepancy | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER HAD ENTERED SAFETY MODE. TECHNICAL SERVICES (TS) WAS CONSULTED BECAUSE THE PACEMAKER WAS UNABLE TO BE INTERROGATED. THIS PACEMAKER REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21263079&pc=LWP | 2124215-2025-04371 | 2025-01-09 05:00 | 2025-01-29 05:00 | Malfunction | Interrogation Problem | ACCOLADE MRI SR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE ATTEMPTED IMPLANT OF THIS PACEMAKER, THE DEVICE COULD NOT BE INTERROGATED. TROUBLESHOOTING WAS PERFORMED AND UPON ATTEMPTING INTERROGATION WITH ANOTHER DEVICE, THE INTERROGATION WAS SUCCESSFUL. IT WAS OPTED TO COMPLETE THE PROCEDURE WITH THE OTHER PACEMAKER. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE HAS NOT BEEN RETURNED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21144734&pc=LWP | 2124215-2025-01065 | 2024-12-21 05:00 | 2025-01-13 05:00 | Malfunction | Interrogation Problem | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT DURING A CLINIC VISIT, THE HEALTH CARE PROFESSIONAL (HCP) CALLED TECHNICAL SERVICES (TS) AND REQUESTED TROUBLESHOOTING ASSISTANCE DUE TO DIFFICULTIES INTERROGATING THE PACEMAKER SYSTEM USING THIS PROGRAMMER. TS WALKED THE HCP THROUGH THE PROGRAMMER TROUBLESHOOTING AND REBOOTING STEPS. AFTER REBOOTING THE PROGRAMMER, THE HCP STATED THAT A MESSAGE WAS DISPLAYED ON THE PROGRAMMER THAT STATED A SOFTWARE UPDATE WAS NEEDED. TS WALKED THE HCP THROUGH THE SOFTWARE UPDATE AND ADVISED THE HCP TO CONTACT THE LOCAL FIELD REPRESENTATIVE IF THE UPDATE FAILS. AT THIS TIME, THE PROGRAMMER AND PACEMAKER REMAIN IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21050401&pc=LWP | 2124215-2024-82681 | 2024-12-18 05:00 | 2024-12-31 05:00 | Malfunction | Interrogation Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS UNABLE TO BE INTERROGATED. TECHNICAL SERVICES (TS) WAS CONSULTED AND TROUBLESHOOTING WAS PERFORMED, HOWEVER, THE ISSUE WAS NOT RESOLVED. THE USE OF A MAGNET TO ELICIT A RESPONSE WAS RECOMMENDED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE.Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS UNABLE TO BE INTERROGATED. TECHNICAL SERVICES (TS) WAS CONSULTED AND TROUBLESHOOTING WAS PERFORMED, HOWEVER, THE ISSUE WAS NOT RESOLVED. THE USE OF A MAGNET TO ELICIT A RESPONSE WAS RECOMMENDED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE. FURTHER INFORMATION WAS RECEIVED THAT THIS DEVICE WAS EXPLANTED DUE TO NORMAL BATTERY DEPLETION (NBD). NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE HAS BEEN RECEIVED FOR ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20932157&pc=LWP | 2124215-2024-79174 | 2024-11-24 05:00 | 2024-12-13 05:00 | Malfunction | Interrogation Problem | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT CONSULT INTERROGATION OF THIS PACEMAKER WAS UNSUCCESSFUL. IT WAS NOTED THAT CONSULT INTERROGATION DIFFICULTIES CAN OCCUR WITH A LOW BATTERY, AND THE PACEMAKER HAD BEEN IMPLANTED FOR NEARLY SIX YEARS. THIS PACEMAKER REMAINS IMPLANTED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21265368&pc=LWP | 2124215-2025-04427 | 2024-08-23 04:00 | 2025-01-29 05:00 | Malfunction | Low impedance | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED LOW OUT OF RANGE IMPEDANCE MEASUREMENTS FOR ITS ASSOCIATED NON BOSTON SCIENTIFIC RIGHT ATRIAL (RA) LEAD. THIS PACEMAKER REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21232547&pc=LWP | 2124215-2025-03511 | 2024-12-26 05:00 | 2025-01-24 05:00 | Malfunction | Low impedance | ACCOLADE EL DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED LOW OUT OF RANGE PACE IMPEDANCE MEASUREMENTS ON THE RIGHT ATRIAL (RA) LEAD RESULTING IN A LEAD SAFETY SWITCH. THE INVOLVED LEAD IS A NON BOSTON SCIENTIFIC PRODUCT. IT WAS NOTED THIS LEAD HAS BEEN IMPLANTED FOR 26 YEARS. THE HEALTH CARE PROFESSIONAL (HCP) STATED THAT REPROGRAMMING WOULD BE PERFORMED AND THE LEAD SAFETY SWITCH FEATURE WOULD BE DEACTIVATED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. AT THIS TIME, THIS DEVICE SYSTEM REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21210380&pc=LWP | 2124215-2025-02956 | 2025-01-08 05:00 | 2025-01-22 05:00 | Malfunction | Low impedance | ACCOLADE EL DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED LOW OUT OF RANGE PACING IMPEDANCE MEASUREMENTS ON THE NON-BOSTON SCIENTIFIC RIGHT ATRIAL (RA) LEAD FOR ONE YEAR. IT WAS NOTED THAT THE PATIENT EXPERIENCED HEART FAILURE SYMPTOMS AND SHOWED UP IN THE EMERGENCY ROOM. CURRENTLY, THIS PRODUCT REMAINS IN SERVICE AND NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21277036&pc=LWP | 2124215-2025-04850 | 2025-01-14 05:00 | 2025-01-30 05:00 | Malfunction | Over-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS OVERSENSING ATRIAL EVENTS ON THE VENTRICULAR CHANNEL. TECHNICAL SERVICES (TS) RECOMMENDED PROGRAMMING CHANGES. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21267223&pc=LWP | 2124215-2025-04522 | 2025-01-07 05:00 | 2025-01-29 05:00 | Malfunction | Over-Sensing | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER ALERTED FOR AN ATRIAL ARRHYTHMIA BURDEN. THE STORED ELECTROGRAMS (EGMS) SHOWED FALSE ATRIAL TACHY RESPONSE (ATR) EPISODES DUE TO FARFIELD R-WAVE OVERSENSING (FFOS). FURTHER REVIEW WAS RECOMMENDED OF PROGRAMMED PARAMETERS TO ENSURE APPROPRIATE ATRIAL ARRHYTHMIA DETECTION. BATTERY LONGEVITY IS NORMAL, AND LEAD MEASUREMENTS ARE STABLE. THE DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21268240&pc=LWP | 2124215-2025-04564 | 2025-01-07 05:00 | 2025-01-29 05:00 | Malfunction | Over-Sensing | ACCOLADE DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED CROSS-TALK OVERSENSING OF ATRIAL EVENTS ON THE VENTRICULAR CHANNEL. TECHNICAL SERVICES (TS) REVIEWED THE EVENT AND PROVIDED REPROGRAMMING RECOMMENDATIONS. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21248479&pc=LWP | 2124215-2025-03889 | 2025-01-02 05:00 | 2025-01-27 05:00 | Malfunction | Over-Sensing | ACCOLADE MRI SR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WITH THIS RIGHT VENTRICULAR (RV) LEAD RECORDED OVERSENSING OF NOISE THAT APPEARS TO BE DUE TO CAUTERY. POSSIBLE LEAD REVISION IS SUSPECTED AS THE IMPLANT DATE IS VERY CLOSE. THE NEXT MORNING THE PRESENTING ELECTROGRAM (EGM) LOOKED CLEAN. THIS PACEMAKER REMAINS IN SERVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21236894&pc=LWP | 2124215-2025-03626 | 2024-12-27 05:00 | 2025-01-25 05:00 | Malfunction | Over-Sensing | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER EXHIBITED FARFIELD OVERSENSING ON THE ATRIAL CHANNEL POST VENTRICULAR PACE. TECHNICAL SERVICES SUGGESTED REPROGRAMMING OPTIONS. THIS DEVICE REMAINS IN SERVICE AND NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21194178&pc=LWP | 2124215-2025-02489 | 2025-01-02 05:00 | 2025-01-20 05:00 | Malfunction | Over-Sensing | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED OVERSENSING ON THE ATRIAL CHANNEL. ADDITIONALLY, THE PATIENT WAS REPORTED TO HAVE THE FLU AND FELT LIGHTHEADED, AND AS A RESULT, FELL. TECHNICAL SERVICES (TS) WAS CONSULTED AND DISCUSSED THAT THE MOST RECENT LEAD TESTS WERE WITHIN NORMAL LIMITS. NO ADVERSE PATIENT EFFECTS WERE REPORTED. AT THIS TIME, THIS DEVICE REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21179687&pc=LWP | 2124215-2025-02094 | 2025-01-02 05:00 | 2025-01-17 05:00 | Malfunction | Over-Sensing | ACCOLADE MRI DR | Dyspnea; Dizziness |  | 1 | Event Description: IT WAS REPORTED THAT THIS PATIENT EXPERIENCED A POSSIBLE OVERSENSING DUE TO ELECTROMAGNETIC INTERFERENCE (EMI). THE PATIENT HAD SYMPTOMS OF DIZZINESS AND FELT LIGHTHEADED. THE DEVICE REMAINS IN SERVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THIS PATIENT EXPERIENCED A POSSIBLE OVERSENSING DUE TO ELECTROMAGNETIC INTERFERENCE (EMI). THE PATIENT HAD SYMPTOMS OF DIZZINESS AND FELT LIGHTHEADED. THE DEVICE REMAINS IN SERVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: CORRECTED FIELDS: H6. IMPACT CODES.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21144158&pc=LWP | 2124215-2025-01116 | 2024-12-25 05:00 | 2025-01-13 05:00 | Malfunction | Over-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THE MINUTE VENTILATION (MV) SENSOR WAS DISABLED BY SIGNAL ARTIFACT MONITOR (SAM). THE ELECTROGRAM (EGM) DETECTED ARTIFACTS LIKELY DUE TO OVERSENSING OF AN ATRIAL ARRHYTHMIA. THE DEVICE REMAINS IMPLANTED. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21144982&pc=LWP | 2124215-2025-01016 | 2024-10-11 04:00 | 2025-01-13 05:00 | Malfunction | Over-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER RECORDED TWO SIGNAL ARTIFACT MONITORING (SAM) EPISODES. HOWEVER, THESE EVENTS ARE NOT VIEWABLE IN THE CURRENT CONSULT TRANSMISSION, WHICH ONLY INCLUDES THE LAST 72 HOURS OF STORED DATA. ACCESSING THE ELECTROGRAM (EGM) WOULD REQUIRE A PRM INTERROGATION. IT WAS NOTED THAT PERFORMING THREE INTERROGATIONS WITHIN A 48-HOUR PERIOD COULD POTENTIALLY TRIGGER SAFETY MODE. THE PATIENT WAS UNDERGOING AN ABLATION PROCEDURE ON THE DAY THE SAM EPISODES WERE RECORDED, MAKING THE PROCEDURE THE SUSPECTED CAUSE. HOWEVER, THE EXACT CAUSE HAS NOT BEEN CONFIRMED AT THIS TIME. THE DEVICE REMAINS IN SERVICE, AND NO ADVERSE EFFECTS ON THE PATIENT WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21124482&pc=LWP | 2124215-2024-83640 | 2024-12-14 05:00 | 2025-01-10 05:00 | Malfunction | Over-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM EXHIBITED FARFIELD SIGNALS ON THE RIGHT ATRIAL (RA) CHANNEL. THESE SIGNALS WERE FALLING ON THE REFRACTORY PERIOD, LEADING TO INAPPROPRIATE ATRIAL TACHY RESPONSE (ATR) MODE SWITCHES. TECHNICAL SERVICES (TS) RECOMMENDED INCREASING THE RIGHT ATRIAL (RA) LEAD INTERVALS. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21106375&pc=LWP | 2124215-2025-00300 | 2024-12-17 05:00 | 2025-01-08 05:00 | Malfunction | Over-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED OVERSENSING ON THE ATRIAL CHANNEL. TECHNICAL SERVICES (TS) WAS CONSULTED AND REPROGRAMMING OPTIONS WERE RECOMMENDED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21045092&pc=LWP | 2124215-2024-82524 | 2024-12-19 05:00 | 2024-12-30 05:00 | Malfunction | Over-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED FARFIELD OVERSENSING ON THE ATRIAL CHANNEL. TECHNICAL SERVICES (TS) WAS CONSULTED, REPROGRAMMING OPTIONS AND SENSITIVITY ADJUSTMENTS WERE RECOMMENDED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED FARFIELD OVERSENSING ON THE ATRIAL CHANNEL. TECHNICAL SERVICES (TS) WAS CONSULTED, AND REPROGRAMMING OPTIONS AND SENSITIVITY ADJUSTMENTS WERE RECOMMENDED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE. ADDITIONAL INFORMATION REPORTED THAT REPROGRAMMING WAS PERFORMED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE. Manufacturer Narrative: CORRECTION: E1,E2,E3: INITIAL REPORTED UPDATED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21032830&pc=LWP | 2124215-2024-82305 | 2024-12-04 05:00 | 2024-12-29 05:00 | Malfunction | Over-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT AN ALERT HAD BEEN TRIGGERED FOR ATRIAL ARRHYTHMIA ALERTAAB) FOR GREATER THAN 0 HOURS IN A 24-HOUR PERIOD. REVIEW OF THE STORED DATA NOTED 10 ATRIAL TACHYCARDIA RESPONSE (ATR) EVENTS. ALL STORED ELECTROGRAMS SHOWED FARFIELD R-WAVE OVERSENSING (FFRWOS) WITH SOME UPPER RATE TRACKING OF ATRIAL TACHYCARDIA AT 120BPM. ROUTINE FOLLOW UP WAS RECOMMENDED. THE DEVICE REMAINS IN SERVICE AND NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS INVESTIGATION WILL BE UPDATED SHOULD FURTHER INFORMATION BE PROVIDED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21028288&pc=LWP | 2124215-2024-82193 | 2024-12-18 05:00 | 2024-12-27 05:00 | Malfunction | Over-Sensing | ACCOLADE MRI EL DR | Presyncope | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS INAPPROPRIATELY SENSING RESPIRATIONS. TECHNICAL SERVICES (TS) WAS CONSULTED AND RECOMMENDED REPROGRAMMING AND CONTINUE TO MONITOR. PATIENT WAS NEAR SYNCOPE WHEN THE HEART RATE DROPPED. THE DEVICE REMAINS IN SERVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21008114&pc=LWP | 2124215-2024-81550 | 2024-09-30 04:00 | 2024-12-23 05:00 | Malfunction | Over-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER RECORDED ATRIAL TACHY RESPONSES (ATR) DUE TO FARFIELD OVERSENSING. INITIALLY, THE ATRS WERE BELIEVED TO BE CAUSED BY NOISE FROM THE RIGHT VENTRICULAR (RV) LEAD DUE TO A LEAD REVISION. HOWEVER, FURTHER INVESTIGATION CONFIRMED THAT THE ISSUE WAS RELATED TO LEFT AGC SENSING, AS NO RV SENSING ISSUES WERE PRESENT, AND THE PATIENT IS NOT DEPENDENT. TECHNICAL SERVICES (TS) RECOMMENDED REPROGRAMMING OPTIONS TO ADDRESS THESE CONCERNS. NO ADVERSE EFFECTS WERE REPORTED FOR THE PATIENT, AND THE DEVICE REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20952754&pc=LWP | 2124215-2024-79807 | 2024-12-04 05:00 | 2024-12-17 05:00 | Malfunction | Over-Sensing | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT A REVIEW OF AN ELECTROCARDIOGRAM (EGM) FOR THE PATIENT WITH THIS PACEMAKER WAS REQUESTED TO DETERMINE WHETHER THE ARRHYTHMIA WAS VENTRICULAR TACHYCARDIA (VT) OR SUPRAVENTRICULAR TACHYCARDIA (SVT). TECHNICAL SERVICES (TS) REVIEWED THE EVENT AND INDICATED THAT THE RHYTHM WAS LIKELY SVT BUT RECOMMENDED OBTAINING A PHYSICIAN'S OPINION FOR CONFIRMATION. TS ALSO NOTED THAT THE PACEMAKER EXHIBITED FAR-FIELD OVERSENSING OF SOME R-WAVES ON THE ATRIAL CHANNEL. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20956416&pc=LWP | 2124215-2024-79997 | 2024-12-10 05:00 | 2024-12-17 05:00 | Malfunction | Over-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS RECENTLY IMPLANTED PACEMAKER EXHIBITED FARFIELD OVERSENSING ON THE RIGHT ATRIAL (RA) CHANNEL. PROGRAMMING ADJUSTMENTS WERE RECOMMENDED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20947954&pc=LWP | 2124215-2024-79633 | 2023-11-02 04:00 | 2024-12-16 05:00 | Malfunction | Over-Sensing | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED INAPPROPRIATE ATRIAL TACHY RESPONSE (ATR) EPISODES DUE TO FAR FIELD OVERSENSING. PROGRAMMING OPTIONS WERE DISCUSSED AND RECOMMENDED. CURRENTLY, THIS PRODUCT REMAINS IN SERVICE AND NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20928467&pc=LWP | 2124215-2024-79084 | 2024-08-10 04:00 | 2024-12-13 05:00 | Malfunction | Over-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THE PACEMAKER INAPPROPRIATELY SWITCHED TO ATRIAL TACHY RESPONSE DUE TO FARFIELD R-WAVE OVERSENSING WHILE THE PATIENT WAS EXPERIENCING A SUPRAVENTRICULAR TACHYCARDIA. TECHNICAL SERVICES DISCUSSED PROGRAMMING OPTIONS. THE PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20912873&pc=LWP | 2124215-2024-78607 | 2024-11-20 05:00 | 2024-12-12 05:00 | Malfunction | Over-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER RECORDED A SIGNAL ARTIFACT MONITORING (SAM) EPISODE, WHICH AUTOMATICALLY DISABLED THE MINUTE VENTILATION (MV) SENSOR. THE HEALTH CARE PROFESSIONAL (HCP) INVOLVED INDICATED THAT THIS PATIENT SUFFERS FROM ATRIAL FIBRILLATION (AF). AT THIS TIME, NO FURTHER INFORMATION IS AVAILABLE. THE DEVICE REMAINS IN SERVICE AND NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20921217&pc=LWP | 2124215-2024-78815 | 2024-08-22 04:00 | 2024-12-12 05:00 | Malfunction | Over-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXHIBITING INAPPROPRIATE ATRIAL TACHY RESPONSE (ATR) DUE TO FARFIELD OVERSENSING. TECHNICAL SERVICES (TS) WAS CONSULTED AND RECOMMENDED REPROGRAMMING. THE DEVICE REMAINS IN SERVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21051817&pc=LWP | 2124215-2024-82753 | 2024-12-12 05:00 | 2025-01-01 05:00 | Malfunction | Over-Sensing; Battery Problem | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS PACEMAKER WAS SUSPECTED TO BE DEPLETING PREMATURELY. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS CONFIRMED THE BATTERY APPEARED TO BE DEPLETING NORMALLY. IT WAS NOTED THAT THIS PACEMAKER HAD EXHIBITED FARFIELD OVERSENSING PREVIOUSLY. BLANKING HAD BEEN ADJUSTED AT THAT TIME. THIS PACEMAKER IS ON AN ADVISORY AND THIS PATIENT WAS SCHEDULED FOR PROPHYLACTIC EXPLANT. ADDITIONAL INFORMATION WAS THEN RECEIVED NOTING THERE IS 7 YEARS REMAINING ON THE BATTERY AND NOT EXHIBITING PREMATURE BATTERY DEPLETION, THIS PATIENT WILL CONTINUE TO BE MONITORED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21114385&pc=LWP | 2124215-2024-83394 | 2024-12-15 05:00 | 2025-01-09 05:00 | Malfunction | Over-Sensing; Data Problem | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAN AN ALERT HAD BEEN GENERATED FOR ATRIAL ARRHYTHMIA BURDEN (AAB) OF AT LEAST 24.0 HOURS IN A 24-HOUR PERIOD. REVIEW OF THE STORED DATA IDENTIFIED THE EPISODE DURATION WAS 80 HOURS AND 41 MINUTES. THE AAB ALERT IS NOW DISABLED. ADDITIONAL INFORMATION WAS RECEIVED THAT MINUTE VENTILATION (MV) SENSOR WAS DISABLED DUE TO A SIGNAL ARTIFACT MONITOR (SAM) EVENT. A PROGRAMMER SESSION IS REQUIRED TO RE-ENABLE MV SENSOR. THE DEVICE REMAINS IN SERVICE AND NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: AS NO FURTHER INFORMATION CONCERNING THIS REPORT IS EXPECTED, OUR INVESTIGATION IS COMPLETE. THIS INVESTIGATION WILL BE UPDATED SHOULD FURTHER INFORMATION BE PROVIDED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21144661&pc=LWP | 2124215-2025-00955 | 2024-12-30 05:00 | 2025-01-13 05:00 | Malfunction | Over-Sensing; Defective Device | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER EXHIBITED FARFIELD OVERSENSING ON THE RIGHT ATRIAL (RA) CHANNEL RESULTING IN INAPPROPRIATE ATRIAL TACHYCARDIA RESPONSE (ATR) EPISODE BEING RECORDED. TECHNICAL SERVICES PROVIDED REPROGRAMMING RECOMMENDATIONS. AT THIS TIME, THIS PACEMAKER SYSTEM REMAINS IN SERVICE AND THERE WERE NO ADVERSE PATIENT EFFECTS REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21270332&pc=LWP | 2124215-2025-04671 | 2025-01-14 05:00 | 2025-01-29 05:00 | Malfunction | Over-Sensing; Failure to Read Input Signal | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED FARFIELD OVERSENSING AND STORED INAPPROPRIATE ATRIAL TACHY RESPONSE (ATR) EPISODES. THE DEVICE WAS PROGRAMMED TO AUTOMATIC GAIN CONTROL (AGC) 0.25 MV AND SMART BLANKING. THE FARFIELD SIGNAL WAS MEASURED AS 0.2 MV ON THE PROGRAMMER, AND P-WAVE AMPLITUDE DAILY MEASUREMENT WAS GREATER THAN 1 MV. UPON REPROGRAMMING TO FIXED SENSING AT NOMINAL VALUES, FARFIELD OVERSENSING WAS RESOLVED. THIS PACEMAKER REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21219780&pc=LWP | 2124215-2025-03191 | 2025-01-13 05:00 | 2025-01-23 05:00 | Malfunction | Over-Sensing; Failure to Read Input Signal | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED VENTRICULAR FAR-FIELD R-WAVE OVERSENSING ON THE ATRIAL ELECTROGRAM (EGM). TECHNICAL SERVICES (TS) WAS CONSULTED FOR FURTHER REVIEW AND PROGRAMMING RECOMMENDATIONS. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PRODUCT REMAINS IN SERVICE. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21197476&pc=LWP | 2124215-2025-02587 | 2025-01-03 05:00 | 2025-01-21 05:00 | Malfunction | Over-Sensing; Failure to Read Input Signal | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM EXHIBITED FARFIELD R-WAVE OVERSENSING AND LOW P-WAVE AMPLITUDES AND THE PATIENT WAS SYNCOPAL, HOWEVER THERE WAS NO EVIDENCE OF UNDERSENSING OR HIGH VENTRICULAR RATES. THE DEVICE WAS OPERATING AS EXPECTED AND OTHER LEAD MEASUREMENTS WERE WITHIN NORMAL LIMITS. THIS PACEMAKER SYSTEM REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21061292&pc=LWP | 2124215-2024-82916 | 2024-12-16 05:00 | 2025-01-02 05:00 | Malfunction | Over-Sensing; Failure to Read Input Signal | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER RECORDED OUT OF RANGE ATRIAL INTRINSIC AMPLITUDE MEASUREMENTS RESULTING IN AN ALERT IN THE REMOTE HOME MONITORING SYSTEM. NO UNDERSENSING WAS OBSERVED AS A RESULT. REVIEW OF STORED ELECTROGRAMS (EGMS) NOTED FARFIELD R-WAVE OVERSENSING ON THE ATRIAL CHANNEL, WHICH, RESULTED IN FREQUENT INAPPROPRIATE ATRIAL TACHY RESPONSE (ATR) MODE SWITCHES. FURTHER REVIEW WAS RECOMMENDED. NO FURTHER ASSISTANCE WAS REQUESTED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE. IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21049284&pc=LWP | 2124215-2024-82683 | 2024-12-18 05:00 | 2024-12-31 05:00 | Malfunction | Over-Sensing; Failure to Read Input Signal | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED OVERSENSING ON THE ATRIAL CHANNEL. LOW AMPLITUDE WAS OBSERVED ON THE RIGHT ATRIAL (RA) LEAD. TECHNICAL SERVICES (TS) RECOMMENDED REPROGRAMMING OPTIONS. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21028284&pc=LWP | 2124215-2024-82182 | 2024-10-29 04:00 | 2024-12-27 05:00 | Malfunction | Over-Sensing; Failure to Read Input Signal | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE HEALTH CARE PROFESSIONAL (HCP) CALLED FOR REVIEW OF A STORED NON-SUSTAINED VENTRICULAR TACHYCARDIA (NSVT) EPISODE FOR THIS PATIENT WITH A PACEMAKER SYSTEM AS THERE WAS A CONCERN THERE WAS T WAVE OVERSENSING. TECHNICAL SERVICES REVIEWED THE DEVICE DATA AND A FEW OLDER EPISODES AND CONFIRMED THERE IS T WAVE OVERSENSING. THE PRESENTING ELECTROGRAM (EGM) SHOWS NORMAL R WAVES AND VERY SMALL T WAVES. THE HCP DID MENTION THAT THE PATIENT HAD A STROKE RIGHT AROUND THIS TIME WHICH SOMETHING MAY HAVE HAPPENED METABOLICALLY. THE PLAN IS TO CONTINUE TO MONITOR. AT THIS TIME, THE DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21151736&pc=LWP | 2124215-2025-01294 | 2024-09-01 04:00 | 2025-01-14 05:00 | Malfunction | Over-Sensing; Failure to Read Input Signal; Under-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED UNDERSENSING OF ATRIAL FIBRILLATION (AF) DUE TO LOW SIGNAL AMPLITUDE MEASUREMENTS (FUNCTIONAL UNDERSENSING). A HEALTH CARE PROFESSIONAL (HCP) ALSO NOTED THAT THE DEVICE EXHIBITED FARFIELD R-WAVE OVERSENSING ON THE ATRIAL CHANNEL 3 MONTHS AGO. ATRIAL SENSITIVITY WAS REPROGRAMMED AT THAT TIME. THE HCP NOTED THAT SENSITIVITY PROGRAMMING WOULD BE REVIEWED IN CLINIC AT THE NEXT SCHEDULED IN CLINIC FOLLOW UP APPOINTMENT. NO FURTHER ASSISTANCE WAS REQUESTED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE. IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20993539&pc=LWP | 2124215-2024-81205 | 2024-11-26 05:00 | 2024-12-20 05:00 | Malfunction | Over-Sensing; Failure to Read Input Signal; Under-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: IT WAS REPORTED THAT THE PACEMAKER ALERTED DUE TO LOW ATRIAL INTRINSIC AMPLITUDE. REVIEW OF THE DEVICE DATA SHOWED THERE TO BE OCCASIONAL FARFIELD OVERSENSING AND ISOLATED UNDERSENSING ON THE ATRIAL THRESHOLD TEST. THE LEAD TREND DATA HAD RECORDED ATRIAL AMPLITUDES OF 0.4 MV TO 5.2 MV AND THE PROGRAMMED ATRIAL SENSITIVITY WAS AUTOMATIC GAIN CONTROL 0.25MV. TECHNICAL SERVICES RECOMMENDED FURTHER REVIEW OF THE PROGRAMMED SETTINGS. THE PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21170043&pc=LWP | 2124215-2025-01794 | 2025-01-02 05:00 | 2025-01-16 05:00 | Malfunction | Over-Sensing; High Capture Threshold | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED A HIGH RIGHT VENTRICULAR (RV) CAPTURE THRESHOLD. TECHNICAL SERVICES (TS) DISCUSSED DEVICE USING HIGH PERCENTAGE OF POWER USAGE. ADDITIONALLY, A REVIEW OF A STORED ATRIAL TACHY RESPONSE (ATR) NOTED POSSIBLE OVERSENSING. AN ANALYSIS OF DEVICE DATA CONFIRMED THIS DEVICE WAS DEPLETING NORMALLY. THIS PACEMAKER REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED A HIGH RIGHT VENTRICULAR (RV) CAPTURE THRESHOLD. TECHNICAL SERVICES (TS) DISCUSSED DEVICE USING HIGH PERCENTAGE OF POWER USAGE. ADDITIONALLY, A REVIEW OF A STORED ATRIAL TACHY RESPONSE (ATR) NOTED POSSIBLE OVERSENSING. AN ANALYSIS OF DEVICE DATA CONFIRMED THIS DEVICE WAS DEPLETING NORMALLY. THE RV LEAD OUTPUTS WERE SET MANUALLY. THIS PACEMAKER REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS REPORT CONTAINS CORRECTIONS TO SECTION H, FIELD H6: IMPACT CODES.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20921397&pc=LWP | 2124215-2024-78886 | 2024-07-24 04:00 | 2024-12-12 05:00 | Malfunction | Over-Sensing; Impedance Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT AN ALERT WAS RECEIVED FROM THIS IMPLANTABLE PACEMAKER, INDICATING THAT A RIGHT ATRIAL (RA) LEAD SAFETY SWITCH (LSS) TO UNIPOLAR SENSING HAD OCCURRED DUE TO OUT-OF-RANGE PACING IMPEDANCE. PRIOR TO THE LSS, A SIGNAL ARTIFACT MONITORING (SAM) EPISODE HAD ALSO BEEN DECLARED DUE TO RA ELECTROMAGNETIC INTERFERENCE (EMI) ARTIFACTS. THE PATIENT WAS SUBSEQUENTLY EVALUATED IN-CLINIC, WHERE PROVOCATIVE MANEUVERS REPRODUCED SOME EVIDENCE OF OVER-SENSED NOISE. THE LEAD THRESHOLD MEASUREMENTS WERE SUITABLE. A LEAD CONDUCTOR FRACTURE WAS SUSPECTED. THE PHYSICIAN ELECTED TO LEAVE THE SYSTEM PROGRAMMED TO THE UNIPOLAR CONFIGURATION AND INCREASE THE AUTOMATIC GAIN CONTROL FEATURE. ANOTHER FOLLOW-UP APPOINTMENT WAS SCHEDULED FOR NEXT MONTH, WHERE THE PHYSICIAN WILL EVALUATE X-RAY IMAGING AND DETERMINE WHETHER A LEAD REVISION WILL BE PERFORMED. AT THIS TIME, THE SYSTEM REMAINS IMPLANTED AND IN-SERVICE. IT IS CURRENTLY UNKNOWN IF THE RA LEAD IS BOSTON SCIENTIFIC PRODUCT. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21285889&pc=LWP | 2124215-2025-05117 | 2025-01-17 05:00 | 2025-01-31 05:00 | Malfunction | Over-Sensing; Pacing Problem | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS PACEMAKER SYSTEM UNDERWENT A SURGICAL PROCEDURE. DURING SURGERY, ABNORMAL FAST PACING WAS OBSERVED. THE HEALTH CARE PROFESSIONAL (HCP) STATED THAT A MAGNET WAS NOT USED DUE TO THE SURGERY BEING DONE IN THE ABDOMINAL AREA. TECHNICAL SERVICES (TS) RECOMMENDED TO CONTACT A LOCAL FIELD REPRESENTATIVE FOR FURTHER EVALUATION OF THIS DEVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21246339&pc=LWP | 2124215-2025-03861 | 2025-01-10 05:00 | 2025-01-27 05:00 | Malfunction | Over-Sensing; Pacing Problem | ACCOLADE MRI EL DR | Dizziness; Asystole | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS OVERSENSING ATRIAL EVENTS ON THE VENTRICULAR CHANNEL WHICH LED TO PACING INHIBITION, AND THE PATIENT WENT TO THE EMERGENCY ROOM DUE TO DIZZINESS. TECHNICAL SERVICES (TS) RECOMMENDED PROGRAMMING CHANGES. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21188308&pc=LWP | 2124215-2025-02321 | 2024-12-23 05:00 | 2025-01-19 05:00 | Malfunction | Over-Sensing; Pacing Problem | ACCOLADE MRI EL DR | Dyspnea | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PATIENT EXPERIENCED SHORTNESS OF BREATH DURING HALL WALKS. UPON REVIEW IT WAS NOTED THAT THIS PACEMAKER RECORDED NOISE OVERSENSING ON THE MINUTE VENTILATION (MV) CHANNEL WHICH DISABLED THE MV FEATURE. TECHNICAL SERVICES (TS) WAS CONSULTED, AND IT WAS ADVISED ON USING THE ACCELEROMETER FEATURE INSTEAD. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21146279&pc=LWP | 2124215-2025-01183 | 2024-12-13 05:00 | 2025-01-13 05:00 | Malfunction | Over-Sensing; Pacing Problem | ACCOLADE MRI SR | Bradycardia | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM WAS EXPERIENCING INTERMITTENT FAR FIELD OVERSENSING OF THE PATIENT T-WAVES. THE PATIENT'S WATCH AND HOLTER MONITOR SHOWED THAT THEIR HEART RATE WAS IN THE 40S AND THE DEVICE WAS PROGRAMMED TO 60 BPM. TECHNICAL SERVICES RECOMMENDED PROGRAMMING OPTIONS TO REDUCE THE OVERSENSING. THIS PACEMAKER SYSTEM REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20984811&pc=LWP | 2124215-2024-80909 | 2024-12-10 05:00 | 2024-12-19 05:00 | Malfunction | Over-Sensing; Pacing Problem | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED FARFIELD R-WAVE OVERSENSING AND STORED PACEMAKER MEDIATED TACHYCARDIA (PMT) EPISODES. THIS PACEMAKER REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20967919&pc=LWP | 2124215-2024-80295 | 2024-12-11 05:00 | 2024-12-18 05:00 | Malfunction | Over-Sensing; Pacing Problem | ACCOLADE MRI DR | Arrhythmia | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED OVERSENSING OF VENTRICULAR SIGNALS ON THE ATRIAL CHANNEL, WHICH TRIGGERED PACEMAKER-MEDIATED TACHYCARDIA (PMT) EPISODES. TECHNICAL SERVICES (TS) WAS CONSULTED AND REPROGRAMMING OPTIONS WERE DISCUSSED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20932158&pc=LWP | 2124215-2024-79175 | 2024-06-16 04:00 | 2024-12-13 05:00 | Malfunction | Over-Sensing; Pacing Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM STORED PACEMAKER MEDIATED TACHYCARDIA (PMT) EPISODES THAT APPEARED TO START WITHOUT PREMATURE VENTRICULAR CONTRACTION (PVC), AND THE PATIENT SEEMED TO HAVE RETROGRADE CONDUCTION. ADDITIONALLY, THERE WAS SOME FARFIELD OVERSENSING OF VENTRICULAR SIGNALS ON THE RIGHT ATRIAL (RA) CHANNEL. TECHNICAL SERVICES (TS) DISCUSSED TROUBLESHOOTING OPTIONS AND PROGRAMMING CONSIDERATIONS. THIS PACEMAKER SYSTEM REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM STORED PACEMAKER MEDIATED TACHYCARDIA (PMT) EPISODES THAT APPEARED TO START WITHOUT PREMATURE VENTRICULAR CONTRACTION (PVC), AND THE PATIENT SEEMED TO HAVE RETROGRADE CONDUCTION. ADDITIONALLY, THERE WAS SOME FARFIELD OVERSENSING OF VENTRICULAR SIGNALS ON THE RIGHT ATRIAL (RA) CHANNEL. TECHNICAL SERVICES (TS) DISCUSSED TROUBLESHOOTING OPTIONS AND PROGRAMMING CONSIDERATIONS. THIS PACEMAKER SYSTEM REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. CORRECTION TO A1: UPDATED TO U.P. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20924548&pc=LWP | 2124215-2024-78962 | 2024-12-04 05:00 | 2024-12-13 05:00 | Malfunction | Over-Sensing; Pacing Problem; Device Sensing Problem | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THERE WERE CONCERNS THAT THERE WAS PACING INHIBITION ON THE RIGHT ATRIAL (RA) AND RIGHT VENTRICULAR (RV) LEAD CHANNELS OF THIS PACEMAKER AS THERE WAS NO PACING MORPHOLOGY DURING IMPLANT. TS ALSO NOTED CONCERNS ABOUT THE LOCATION AS THERE APPEARED TO BE DEVICE SENSING ISSUES, SUCH AS OVERSENSING. THE PATIENT'S WOUND WAS ALREADY CLOSED AT THE TIME OF THE CALL. THUS, NO TROUBLESHOOTING WAS DONE. THE DEVICE REMAINS IN USE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21148332&pc=LWP | 2124215-2025-01225 | 2018-01-01 05:00 | 2025-01-14 05:00 | Malfunction | Over-Sensing; Pacing Problem; Under-Sensing; Battery Problem | ACCOLADE EL DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT ATRIAL UNDER SENSING WAS OBSERVED ON THIS PACEMAKER SYSTEM. TECHNICAL SERVICES (TS) ADVISED THIS IS FUNCTIONAL UNDER SENSING. THE PATIENT IS PACEMAKER DEPENDENT. ADDITIONAL REVIEW FOUND UNDER SENSING HAD RESULTED IN FALSE PACEMAKER MEDIATED TACHYCARDIA EPISODES TO BE STORED. TS ADVISED THERE WAS EVIDENCE OF SOME ATRIAL FAR-FIELD OVERSENSING. TS FOUND THERE IS INCREASED POWER CONSUMPTION FOR PERSISTENT HIGH-RATE ATRIAL SENSING AND THE RIGHT VENTRICULAR PACING AT A HIGHER AVERAGE PACING RATE. THE PATIENT IS ON AN ADVISORY FOR ACCOLADE HIGH BATTERY IMPEDANCE AND IS EXPECTED TO HAVE A CARDIOVERSION AND A DEVICE CHANGE IN THE FUTURE. TS NOTED THE DEVICE HAS BEEN RECENTLY REPROGRAMMED. THE PACEMAKER SYSTEM REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21242093&pc=LWP | 2124215-2025-03747 | 2025-01-13 05:00 | 2025-01-27 05:00 | Malfunction | Over-Sensing; Telemetry Discrepancy | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXPERIENCED RADIO FREQUENCY (RF) OVERUSE. TECHNICAL SERVICES (TS) WAS CONSULTED AND NOTED INTERMITTENT MYOPOTENTIAL RIGHT VENTRICULAR (RV) LEAD OVERSENSING. THE DEVICE REMAINS IN SERVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21257408&pc=LWP | 2124215-2025-04232 | 2025-01-09 05:00 | 2025-01-28 05:00 | Malfunction | Over-Sensing; Under-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED FAR-FIELD OVERSENSING OF THE R-WAVE ON THE ATRIAL CHANNEL, CAUSING INAPPROPRIATE ATRIAL TACHY RESPONSE (ATR). ADDITIONALLY, UNDERSENSING EPISODES RELATED TO THE PATIENT'S HISTORY OF JUNCTIONAL RHYTHM WERE NOTED. TECHNICAL SERVICES (TS) PROVIDED REPROGRAMMING OPTIONS. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21214626&pc=LWP | 2124215-2025-03060 | 2025-01-05 05:00 | 2025-01-23 05:00 | Malfunction | Over-Sensing; Under-Sensing | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE HEALTH CARE PROFESSIONAL (HCP) CALLED TECHNICAL SERVICES (TS) AND REQUESTED REVIEW OF THIS PACEMAKER STORED ATRIAL TACHY RESPONSE (ATR) EPISODES. IT WAS NOTED THAT THE PATIENT WAS IN AN ATRIAL FIBRILLATION (AF) WITH RAPID VENTRICULAR RESPONSE (RAPID VENTRICULAR RESPONSE (RVR). UPON REVIEW, THE PRESENTING ELECTROGRAM (EGM) SHOWED THE DEVICE EXHIBIT P WAVE UNDERSENSING AND FARFIELD OVERSENSING ON THE RIGHT ATRIAL (RA) LEAD CHANNEL. TS DISCUSSED REPROGRAMMING OPTIONS WITH THE HCP AND RECOMMENDED FOR XRAY IMAGES TO BE SCHEDULED FOR FURTHER EVALUATION. NO ADVERSE PATIENT EFFECTS WERE REPORTED. AT THIS TIME, THE PACEMAKER AND RA LEAD REMAIN IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21214416&pc=LWP | 2124215-2025-03091 | 2025-01-05 05:00 | 2025-01-22 05:00 | Malfunction | Over-Sensing; Under-Sensing | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM EXHIBITED P-WAVE UNDERSENSING AS WELL AS FAR FIELD OVERSENSING ON THE RIGHT ATRIAL (RA) LEAD CHANNEL. TECHNICAL SERVICES (TS) WAS CONSULTED AND ADVISED ON FURTHER IN OFFICE REVIEW. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS SYSTEM REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21214414&pc=LWP | 2124215-2025-03092 | 2024-12-26 05:00 | 2025-01-22 05:00 | Malfunction | Pacing Problem | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT PRIOR TO THE IMPLANT OF THIS PACEMAKER, THE DEVICE WAS PROGRAMMED TO PACE AT 60 BEATS PER MINUTE, HOWEVER IT WAS NOTED THAT THE PACING WAS AT 100 BEATS PER MINUTE. IT WAS OPTED TO COMPLETE THE PROCEDURE WITH ANOTHER DEVICE. TECHNICAL SERVICES (TS) WAS CONSULTED, AND THE CAUSE FOR THE PACING RATE WAS NOT DETERMINED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE HAS NOT BEEN RETURNED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21124478&pc=LWP | 2124215-2025-00752 | 2024-11-11 05:00 | 2025-01-10 05:00 | Malfunction | Pacing Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE HEALTH CARE PROFESSIONAL (HCP) REQUESTED A REVIEW OF AN EPISODE ON THIS PACEMAKER. TECHNICAL SERVICES (TS) REVIEWED THE EPISODE AND DETERMINED IT WAS ATRIAL TACHYCARDIA THAT SELF-TERMINATED, FOLLOWED BY A VISIBLE BUT UNSENSED FARFIELD SIGNAL ON THE PRESENTING ELECTROGRAM (EGM). TS SPECULATED WHETHER THE RHYTHM COULD BE RELATED TO PATIENT MOVEMENT OR PHYSICAL ACTIVITY. AFTER AN IN-CLINIC VISIT, NO SENSOR RATE INCREASES OR FARFIELD SIGNALS WERE NOTED DURING A POSITIONING TEST. THE MINUTE VENTILATION (MV) AND ACCELEROMETER SENSOR RATES VARIED WITH SPECIFIC MOTIONS, AND THE HEART RATE DID NOT INCREASE AS EXPECTED DURING PHYSICAL ACTIVITY SESSIONS. TS PROVIDED REPROGRAMMING RECOMMENDATIONS. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE. Manufacturer Narrative: THIS PRODUCT HAS NOT BEEN RETURNED TO BOSTON SCIENTIFIC, AND AS A RESULT, LABORATORY ANALYSIS COULD NOT BE CONDUCTED. THE ASSOCIATED INVESTIGATION DETERMINED THAT THIS DEVICE APPEARED TO EXHIBIT AN INCREASE IN THE PACING RATE AS A RESULT OF THE MINUTE VENTILATION (MV) SENSOR RESPONSE TO INSTANCES WHERE THE MINUTE VENTILATION (MV) SENSOR DID NOT RESPOND AS EXPECTED AT HIGHER RATES. PLEASE REFER TO THE DESCRIPTION FOR MORE INFORMATION REGARDING THE SPECIFIC CIRCUMSTANCES OF THIS EVENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21126750&pc=LWP | 2124215-2025-00822 | 2024-10-29 04:00 | 2025-01-10 05:00 | Malfunction | Pacing Problem | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PATIENT UNDERWENT A TREADMILL TEST DUE TO AN INAPPROPRIATE HEART RATE PER THEIR APPLE WATCH. IT WAS NOTED THE PATIENT'S HEART RATE WOULD SPIKE WITH MOVEMENT OF THEIR LEFT ARM FOLLOWING EXERCISE, ALTHOUGH THIS WAS NOT REPRODUCIBLE IN CLINIC. THE PATIENT ALSO COMPLAINED OF A SPIKING SENSATION IN THEIR CHEST THEY BELIEVE IS DUE TO THE MINUTE VENTILATION (MV) SENSOR, WHICH THEY NOTED FEELS DIFFERENT FROM THE RATTLING SENSATION WHEN THEY GO INTO PAROXYSMAL ATRIAL FLUTTER. TECHNICAL SERVICES (TS) WAS CONSULTED AND REQUESTED THE TREADMILL DATA AND RAW MV SENSOR DATA. TS ALSO NOTED THE CLINIC CAN CONSIDER MAGNET-TRIGGERED ELECTROGRAM (EGM) TO DIFFERENTIATE PAROXYSMAL ATRIAL FLUTTER VERSUS INAPPROPRIATE SENSOR-DRIVEN PACING. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PRODUCT REMAINS IN SERVICE. Manufacturer Narrative: THIS PRODUCT HAS NOT BEEN RETURNED TO BOSTON SCIENTIFIC, AND AS A RESULT, LABORATORY ANALYSIS COULD NOT BE CONDUCTED. THE ASSOCIATED INVESTIGATION DETERMINED THIS DEVICE APPEARED TO EXHIBIT AN INCREASE IN THE PACING RATE AS A RESULT OF THE MINUTE VENTILATION (MV) SENSOR RESPONSE TO UPPER BODY PHYSICAL MOVEMENTS/NON-RESPIRATORY SIGNALS. PLEASE REFER TO THE EVENT DESCRIPTION FOR MORE INFORMATION REGARDING THE SPECIFIC CIRCUMSTANCES OF THIS EVENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20984141&pc=LWP | 2124215-2024-80880 | 2024-10-08 04:00 | 2024-12-19 05:00 | Malfunction | Pacing Problem | ACCOLADE MRI DR | Bradycardia; Dyspnea; Discomfort; Unspecified Heart Problem | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS PACEMAKER DEVICE CALLED TECHNICAL SERVICES (TS) AND NOTED WHEN THEIR LEFT ARM IS RAISED IN PHYSICAL THERAPY, THEIR HEART RATE INCREASES TO 100 BEATS PER MINUTE (BPM) EVEN WITHOUT EXERTION. TS DISCUSSED POSSIBLE CAUSES WITH THE PATIENT. TS RECOMMENDED THE PATIENT CONSULTS WITH THE DEVICE TREATING PHYSICIAN FOR FURTHER EVALUATION OF THE PACEMAKER. AT THIS TIME, THE PACEMAKER REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. SUBSEQUENT ADDITIONAL INFORMATION WAS PROVIDED THAT REPORTED THAT DURING AN IN-CLINIC VISIT, THE PACEMAKER WAS INTERROGATED AND EVALUATED. THE PHYSICIAN DETERMINED THAT THE MINUTE VENTILATION (MV) FEATURE WAS THE CAUSING THE PATIENT HEART RATE TO BE INCREASED WITH ARM EXERCISES DUE TO DEVICE MOVEMENT IN THE POCKET. THE PHYSICIAN REPROGRAMMED THE DEVICE SETTINGS AND RESOLVED THE ISSUE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THE DEVICE REMAINS IN SERVICE. ADDITIONAL INFORMATION WAS RECEIVED THAT REPORTED THAT THE PATIENT CALLED TECHNICAL SERVICES (TS) AND REPORTED EXPERIENCING CHEST DISCOMFORT, SHORTNESS OF BREATH (SOB) AND BRADYCARDIA SYMPTOMS. THE PATIENT REPORTED HISTORY OF ATRIAL FIBRILLATION (AF) AND ABNORMAL HEART BEATS. TS REVIEWED THE STORED DEVICE DATA LOGS AND DISCUSSED WITH THE PATIENT THAT THE STORED DEVICE DATA SHOWED ARRHYTHMIAS AND ECTOPIC BEATS DEPENDING ON RHYTHM AND DEVICE PROGRAMMING. TS ADVISED THE PATIENT TO FOLLOW UP WITH THE DEVICE TREATING CLINICIAN. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THE DEVICE REMAINS IN SERVICE. Manufacturer Narrative: THIS PRODUCT HAS NOT BEEN RETURNED TO BOSTON SCIENTIFIC, AND AS A RESULT, LABORATORY ANALYSIS COULD NOT BE CONDUCTED. THE ASSOCIATED INVESTIGATION DETERMINED THAT THIS DEVICE APPEARED TO EXHIBIT AN INCREASE IN THE PACING RATE AS A RESULT OF THE MINUTE VENTILATION (MV) SENSOR RESPONSE TO INSTANCES WHERE THE MINUTE VENTILATION (MV) SENSOR DID NOT RESPOND AS EXPECTED AT HIGHER RATES. PLEASE REFER TO THE DESCRIPTION FOR MORE INFORMATION REGARDING THE SPECIFIC CIRCUMSTANCES OF THIS EVENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21145877&pc=LWP | 2124215-2025-01151 | 2024-12-22 05:00 | 2025-01-13 05:00 | Malfunction | Pacing Problem; Failure to Read Input Signal; Under-Sensing | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: IT WAS REPORTED THAT AN ALERT WAS GENERATED FOR ATRIAL INTRINSIC AMPLITUDE OUT OF RANGE. REVIEW OF THE TRENDED DATA SHOWED ATRIAL AMPLITUDES RANGED FROM 0.4MV TO 4.8MV. PRESENTING ELECTROGRAM AND RYTHMIQ EPISODES SHOWED UNDERSENSED ATRIAL BEATS FOLLOWED BY ATRIAL PACING. ATRIAL SENSITIVITY IS PROGRAMMED TO FIXED AT 0.75MV. FURTHER EVALUATION SHOULD BE CONSIDERED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS REPORT IS BEING SUBMITTED WITH CORRECTED EVENT DESCRIPTION.Event Description: LATITUDETM ALERT TRANSMISSION SUMMARY REPORT PATIENT DETAILS / PRIORITY / ALERT REASON PATIENT NAME: (B)(6). DEVICE: L311 ACCOLADE MRI/(B)(6). PRIORITY: 2. FURTHER REVIEW RECOMMENDED. ALERT: ATRIAL INTRINSIC AMPLITUDE OUT OF RANGE. SUMMARY: ATRIAL INTRINSIC AMPLITUDE OUT OF RANGE. UNDERSENSED ATRIAL BEAT FOLLOWED BY AP NOTED ON PRESENTING EGM AND RYTHMIQ EPISODES. TREND SHOWS LOWEST ATRIAL INTRINSIC AMPLITUDE OF 0.4MV, HIGHEST OF 4.8MV. ATRIAL SENSITIVITY FIXED AT 0.75MV. CONSIDER FURTHER REVIEW. BATTERY STATUS OK. OTHER LEAD MEASUREMENTS SATISFACTORY. FOR FURTHER QUESTIONS, CONTACT RHYTHMCARE ANZ TECHNICAL SERVICES AT 1.800.245.559 OR RHYTHMCARE_ANZ@BSCI.COM. (B)(4) REMOTE CLINICAL SPECIALIST RHYTHMCARETM ANZ. Manufacturer Narrative: AS NO FURTHER INFORMATION CONCERNING THIS REPORT IS EXPECTED, OUR INVESTIGATION IS COMPLETE. THIS INVESTIGATION WILL BE UPDATED SHOULD FURTHER INFORMATION BE PROVIDED. Manufacturer Narrative: AS NO FURTHER INFORMATION CONCERNING THIS REPORT IS EXPECTED, OUR INVESTIGATION IS COMPLETE. THIS INVESTIGATION WILL BE UPDATED SHOULD FURTHER INFORMATION BE PROVIDED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21123653&pc=LWP | 2124215-2025-00638 | 2024-12-11 05:00 | 2025-01-10 05:00 | Malfunction | Pacing Problem; Low impedance | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM EXHIBITED LOW OUT OF RANGE PACING IMPEDANCES MEASURING BELOW 200 OHMS ON THE RIGHT ATRIAL (RA) LEAD CHANNEL WHICH LED TO A LEAD SAFETY SWITCH. UPON IN OFFICE REVIEW, A HIGH PACING RATE WAS NOTED, AND REPROGRAMMING OF THE MINUTE VENTILATION SENSOR VECTORS WAS PERFORMED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS SYSTEM REMAINS IN SERVICE.Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM EXHIBITED LOW OUT OF RANGE PACING IMPEDANCES MEASURING BELOW 200 OHMS ON THE RIGHT ATRIAL (RA) LEAD CHANNEL WHICH LED TO A LEAD SAFETY SWITCH. UPON IN OFFICE REVIEW, A HIGH PACING RATE WAS NOTED, AND REPROGRAMMING OF THE MINUTE VENTILATION SENSOR VECTORS WAS PERFORMED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS SYSTEM REMAINS IN SERVICE. Manufacturer Narrative: CORRECTION TO SECTION H DEVICE MANUFACTURERS ONLY, FIELD H6 DEVICE CODES.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21275670&pc=LWP | 2124215-2025-04810 | 2024-12-23 05:00 | 2025-01-30 05:00 | Malfunction | Pacing Problem; Under-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED UNDERSENSING OF PREMATURE VENTRICULAR CONTRACTION (PVC) EVENTS. ADDITIONALLY, THIS DEVICE HAD STORED EPISODES OF PACEMAKER MEDIATED TACHYCARDIA (PMT). THE PMT ALGORITHM SUCCESSFULLY TERMINATED THESE EPISODES. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21222264&pc=LWP | 2124215-2025-03204 | 2024-05-09 04:00 | 2025-01-23 05:00 | Malfunction | Pacing Problem; Under-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THERE WAS AN ATRIAL FIBRILLATION (AF) IN PROGRESS WITH OCCASIONAL UNDERSENSING OF AF WITH INAPPROPRIATE ATRIAL PACING NOTED ON THE STORED ELECTROGRAM (EGM) AS WELL AS FREQUENT MODE SWITCHES. THE ATRIAL SENSITIVITY WAS REPROGRAMMED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THE DEVICE REMAINS IMPLANTED.Event Description: IT WAS REPORTED THAT THERE WAS AN ATRIAL FIBRILLATION (AF) IN PROGRESS WITH OCCASIONAL UNDERSENSING OF AF WITH INAPPROPRIATE ATRIAL PACING NOTED ON THE STORED ELECTROGRAM (EGM) AS WELL AS FREQUENT MODE SWITCHES. THE ATRIAL SENSITIVITY WAS REPROGRAMMED. ADDITIONAL INFORMATION NOTED THAT WHEN REVIEWING THE EGM VIA REMOTE MONITORING THE CAUSE OF THE MODE SWITCHED WAS DUE TO ATRIAL FIBRILLATION AND THERE WAS EVIDENCE THAT SOME OF THE AFIB SIGNALS WERE TOO SMALL TO BE SENSED EVEN AT THE MAXIMUM DEVICE SENSITIVITY. NO FURTHER PROGRAMMING CHANGES WERE MADE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THE DEVICE REMAINS IMPLANTED.Event Description: IT WAS REPORTED THAT THERE WAS AN ATRIAL FIBRILLATION (AF) IN PROGRESS WITH OCCASIONAL UNDERSENSING OF AF WITH INAPPROPRIATE ATRIAL PACING NOTED ON THE STORED ELECTROGRAM (EGM) AS WELL AS FREQUENT MODE SWITCHES. THE ATRIAL SENSITIVITY WAS REPROGRAMMED. ADDITIONAL INFORMATION NOTED THAT WHEN REVIEWING THE EGM VIA REMOTE MONITORING THE CAUSE OF THE MODE SWITCHED WAS DUE TO ATRIAL FIBRILLATION AND THERE WAS EVIDENCE THAT SOME OF THE AFIB SIGNALS WERE TOO SMALL TO BE SENSED EVEN AT THE MAXIMUM DEVICE SENSITIVITY. NO FURTHER PROGRAMMING CHANGES WERE MADE. ADDITIONAL INFORMATION NOTED THAT A REVIEW OF PRESENTING ELECTROGRAM (EGM) SHOWED VENTRICULAR UNDER SENSING RESULTING IN INAPPROPRIATE VENTRICULAR PACING. IT WAS DISCUSSED TO CONSIDER INCREASING THE RIGHT VENTRICULAR SENSITIVITY AT THE NEXT IN CLINIC REVIEW. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THE DEVICE REMAINS IMPLANTED. Manufacturer Narrative: THIS REPORT IS BEING FILED TO UPDATE THE DESCRIBE EVENT OR PROBLEM FIELD. Manufacturer Narrative: THIS REPORT IS BEING FILED TO UPDATE THE DESCRIBE EVENT OR PROBLEM FIELD.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21144795&pc=LWP | 2124215-2024-83820 | 2024-12-17 05:00 | 2025-01-13 05:00 | Malfunction | Pacing Problem; Under-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED UNDERSENSING OF R-WAVES. THE SIGNALS WERE REDUCED, BUT WERE STILL ABOVE THE PROGRAMMED SENSITIVITY OF 0.7MV. THE UNDERSENSING CAUSED INAPPROPRIATE MODE SWITCHING FROM AAI TO DDD MODE, RESULTING IN PACING BEING DELIVERED WHEN NOT NEEDED. TECHNICAL SERVICES DISCUSSED HOW THE AUTOMATIC GAIN CONTROL (AGC) FEATURE WORKS, NOTING THAT WITH ONLY OCCASIONAL SMALL AMPLITUDE BEATS, THE AGC DOES NOT STEP DOWN TO THE SENSING FLOOR BEFORE LARGER AMPLITUDES ARE OBSERVED. TECHNICAL SERVICES ALSO DISCUSSED THAT IT MAY BE MORE CLINICALLY APPROPRIATE TO CONSIDER AV SEARCH+ RATHER THAN RYTHMIQ TO MINIMIZE UNNECESSARY RV PACING. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THE DEVICE REMAINS IMPLANTED AND IN SERVICE. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21024644&pc=LWP | 2124215-2024-82094 | 2023-10-11 04:00 | 2024-12-27 05:00 | Malfunction | Pocket Stimulation | ACCOLADE MRI EL DR | Device Overstimulation of Tissue | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) WAS A CASE OF AN ATTEMPTED IMPLANT DUE TO CONSISTENT REPRODUCIBLE POCKET STIMULATION WITH ATRIAL PACE (AP) AND VENTRICULAR PACE (VP) THAT WAS PRESENT UPON CONNECTING TO THE DEVICE. HOWEVER, WHEN CONNECTED TO A NEW PROGRAMMER, THE STIMULATION WAS NOT PRESENT. LEAD IMPEDANCES WERE STABLE FROM THE PACING SYSTEM ANALYZER (PSA) WITH NO VISIBLE DAMAGE ON THE LEADS. THIS DEVICE WAS REPLACED AND NEVER IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: UPON RECEIPT AT OUR POST MARKET QUALITY ASSURANCE LABORATORY, THIS DEVICE WAS THOROUGHLY INSPECTED AND ANALYZED. THE MANUFACTURING DEFICIENCY INVESTIGATION CONCLUSION CODE WAS SELECTED BASED ON ANALYSIS EVIDENCE OF A HOLE IN THE BATTERY LINER/COVER, ALLOWING CONTACT BETWEEN THE BATTERY AND THE ELECTRICALLY ACTIVE PG CASE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21256489&pc=LWP | 2124215-2025-04173 | 2025-01-27 05:00 | 2025-01-28 05:00 | Malfunction | Premature Discharge of Battery | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT A REQUEST FOR BATTERY ANALYSIS WAS NEEDED WHEN IT COMES TO THIS DEVICE. TECHNICAL SERVICES (TS) REVIEWED THE CASE AND CONFIRMED THAT THE DEVICE POWER CONSUMPTION WAS INCREASING IN A GRADUAL FASHION AND THE DEVICE WAS EXHIBITING PREMATURE BATTERY DEPLETION. TS RECOMMENDED THE DEVICE BE EITHER REPLACED OR HAVE THE BATTERY STATUS RE-EVALUATED IN NINETY DAYS. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THE DEVICE REMAINS IMPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21204249&pc=LWP | 2124215-2025-02802 | 2024-12-25 05:00 | 2025-01-21 05:00 | Malfunction | Premature Discharge of Battery | ACCOLADE DR | No Clinical Signs, Symptoms or Conditions | N970003/S167 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER EXHIBITED PREMATURE BATTERY DEPLETION (PBD). THIS DEVICE IS SCHEDULED FOR AN EARLY REPLACEMENT AND TO DATE, THE DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21169227&pc=LWP | 2124215-2025-01763 | 2024-12-19 05:00 | 2025-01-16 05:00 | Malfunction | Premature Discharge of Battery | ACCOLADE EL DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER DEVICE WAS SUSPECTED OF PREMATURE BATTERY DEPLETION (PBD). DATA ANALYSIS CONFIRMED THE BATTERY APPEARED TO BE DEPLETING MORE QUICKLY THAN EXPECTED. TECHNICAL SERVICES CONSULTANT RECOMMENDED REPLACEMENT. THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21170044&pc=LWP | 2124215-2025-01789 | 2024-11-09 05:00 | 2025-01-16 05:00 | Malfunction | Premature Discharge of Battery | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER RECORDED A CODE 1003 INDICATIVE OF BATTERY VOLTAGE TOO LOW FOR THE PROJECTED REMAINING CAPACITY. THE DEVICE BATTERY WAS SUSPECTED TO BE DEPLETING PREMATURELY. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS CONFIRMED THE BATTERY APPEARED TO BE DEPLETING MORE QUICKLY THAN EXPECTED. DEVICE REPLACEMENT WAS RECOMMENDED, AND A REPLACEMENT INTERVAL WAS DISCUSSED. THE SPECIFIC DEVICE IMPLANT DATE IS UNKNOWN AT THIS TIME. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE. IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.Event Description: IT WAS REPORTED THAT THIS PACEMAKER RECORDED A CODE 1003 INDICATIVE OF BATTERY VOLTAGE TOO LOW FOR THE PROJECTED REMAINING CAPACITY. THE DEVICE BATTERY WAS SUSPECTED TO BE DEPLETING PREMATURELY. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS CONFIRMED THE BATTERY APPEARED TO BE DEPLETING MORE QUICKLY THAN EXPECTED. DEVICE REPLACEMENT WAS RECOMMENDED, AND A REPLACEMENT INTERVAL WAS DISCUSSED. THE SPECIFIC DEVICE IMPLANT DATE IS UNKNOWN AT THIS TIME. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED TO OBTAIN THE DEVICE IMPLANT DATE AND PLAN OF ACTION FOR THIS DEVICE, HOWEVER, NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. ADDITIONAL INFORMATION WAS RECEIVED THAT SURGICAL INTERVENTION WAS UNDERTAKEN. THE DEVICE WAS EXPLANTED AND REPLACED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THE PRODUCT IS EXPECTED TO BE RETURNED FOR ANALYSIS. THIS REPORT WILL BE UPDATED UPON RETURN AND COMPLETION OF ANALYSIS. Manufacturer Narrative: ADDITIONAL INFORMATION HAS BEEN REQUESTED TO OBTAIN THE DEVICE IMPLANT DATE AND PLAN OF ACTION FOR THIS DEVICE, HOWEVER, NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: ADDITIONAL INFORMATION HAS BEEN REQUESTED TO OBTAIN THE DEVICE IMPLANT DATE AND PLAN OF ACTION FOR THIS DEVICE, HOWEVER, NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21160481&pc=LWP | 2124215-2025-01518 | 2024-12-27 05:00 | 2025-01-15 05:00 | Malfunction | Premature Discharge of Battery | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS PACEMAKER WAS SUSPECTED TO BE DEPLETING PREMATURELY. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS CONFIRMED THE BATTERY WAS DEPLETING MORE QUICKLY THAN EXPECTED. A REPLACEMENT INTERVAL WINDOW WAS PROVIDED FOR THIS PACEMAKER. THIS PACEMAKER REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21161803&pc=LWP | 2124215-2025-01584 | 2024-12-17 05:00 | 2025-01-15 05:00 | Malfunction | Premature Discharge of Battery | ACCOLADE MRI DR | Atrial Fibrillation | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER DEVICE WAS EXHIBITING FASTER BATTERY DRAIN. BOSTON SCIENTIFIC TECHNICAL SERVICES (TS) DISCUSSED THAT THE PATIENT HAS PERSISTENT ATRIAL FIBRILLATION (AF) AND THE CONSUMPTION WAS 51 MICROWATTS. ADDITIONALLY, TS SUGGESTED PROGRAMMING THE DEVICE TO VVIR WITH ATRIAL SENSING OFF TO CONSERVE BATTERY. THE DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21123665&pc=LWP | 2124215-2025-00668 | 2024-12-11 05:00 | 2025-01-10 05:00 | Malfunction | Premature Discharge of Battery | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS IMPLANTABLE PACEMAKER WENT DOWN TO 1 YEAR LONGEVITY WHEN HAS ONLY BEEN IN SERVICE FOR ABOUT 8 MONTHS. A REQUEST WAS MADE TO HAVE DATA FROM THE DEVICE ANALYZED. DATA ANALYSIS CONFIRMED THE POWER CONSUMPTION IS ABNORMALLY HIGH AND STABLE. THERE ARE NO FAULTS, AND THE LEAD DIAGNOSTIC APPEARED TO BE NORMAL. DEVICE REPLACEMENT IS STILL RECOMMENDED SINCE ROOT CAUSE AND FUTURE PERFORMANCE OF THE DEVICE ARE UNKNOWN. THE DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21031024&pc=LWP | 2124215-2024-82252 | 2024-10-06 04:00 | 2024-12-27 05:00 | Malfunction | Premature Discharge of Battery | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER RECORDED A CODE 1003 INDICATIVE OF BATTERY VOLTAGE TOO LOW FOR THE PROJECTED REMAINING CAPACITY. DATA ANALYSIS SHOWED THE BATTERY APPEARED TO BE DEPLETING MORE QUICKLY THAN EXPECTED. DEVICE REPLACEMENT WAS RECOMMENDED. AT THIS TIME, THIS PACEMAKER REMAINS IN SERVICE, AND THERE WERE NO ADVERSE PATIENT EFFECTS REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21021147&pc=LWP | 2124215-2024-81980 | 2024-12-10 05:00 | 2024-12-26 05:00 | Malfunction | Premature Discharge of Battery | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS PACEMAKER WAS SUSPECTED TO BE DEPLETING PREMATURELY. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED AND DATA ANALYSIS CONFIRMED THE BATTERY APPEARED TO BE DEPLETING MORE QUICKLY THAN EXPECTED. DEVICE REPLACEMENT WAS RECOMMENDED. THE PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20980942&pc=LWP | 2124215-2024-80765 | 2024-12-09 05:00 | 2024-12-19 05:00 | Malfunction | Premature Discharge of Battery | ACCOLADE EL DR | No Clinical Signs, Symptoms or Conditions | N970003/S167 | 1 | Event Description: IT WAS REPORTED THAT THIS DEVICE RECORDED A CODE 1003 INDICATIVE OF BATTERY VOLTAGE TOO LOW FOR THE PROJECTED REMAINING CAPACITY. THE BATTERY LONGEVITY PROJECTION REDUCED BY TWO YEARS IN APPROXIMATELY NINE MONTHS. IT WAS REPORTED THAT A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS SHOWED THE BATTERY APPEARED TO BE DEPLETING MORE QUICKLY THAN EXPECTED. DEVICE REPLACEMENT WAS RECOMMENDED. THIS DEVICE WAS EXPLANTED AND RETURNED FOR ANALYSIS. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20968380&pc=LWP | 2124215-2024-80342 | 2024-12-09 05:00 | 2024-12-18 05:00 | Malfunction | Premature Discharge of Battery | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THERE WERE CONCERNS THIS PACEMAKER EXHIBITED PREMATURE BATTERY DEPLETION (PBD). THE DEVICE REMAINS IN USE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS INVESTIGATION WILL BE UPDATED WHEN FURTHER INFORMATION BECOMES AVAILABLE. ADDITIONAL INFORMATION RECEIVED INDICATES THAT TECHNICAL SERVICES (TS) REVIEWED THE REPORTS AND FOUND THAT THE DEVICE WAS PREMATURELY DEPLETING. TS ADVISED DEVICE REPLACEMENT OR RE-EVALUATING THE BATTERY IN 90 DAYS. THE DEVICE REMAINS IN USE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20952751&pc=LWP | 2124215-2024-79806 | 2024-12-11 05:00 | 2024-12-17 05:00 | Malfunction | Premature Discharge of Battery | ACCOLADE MRI SR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS PACEMAKER WAS SUSPECTED TO BE DEPLETING PREMATURELY. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS CONFIRMED THE BATTERY APPEARED TO BE DEPLETING MORE QUICKLY THAN EXPECTED. A REPLACEMENT INTERVAL WINDOW WAS PROVIDED FOR THIS PACEMAKER. THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20940287&pc=LWP | 2124215-2024-79366 | 2024-12-09 05:00 | 2024-12-16 05:00 | Malfunction | Premature Discharge of Battery | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS PACEMAKER WAS SUSPECTED TO BE DEPLETING PREMATURELY. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS CONFIRMED THE BATTERY APPEARED TO BE DEPLETING MORE QUICKLY THAN EXPECTED. A REPLACEMENT INTERVAL WINDOW WAS PROVIDED FOR THIS PACEMAKER. THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20908280&pc=LWP | 2124215-2024-78461 | 2024-12-03 05:00 | 2024-12-11 05:00 | Malfunction | Premature Discharge of Battery | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS SUSPECTED TO BE DEPLETING PREMATURELY. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS CONFIRMED THE BATTERY APPEARED TO BE DEPLETING MORE QUICKLY THAN EXPECTED. DEVICE REPLACEMENT WAS RECOMMENDED, AND A REPLACEMENT INTERVAL WAS DISCUSSED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21276064&pc=LWP | 2124215-2025-04788 | 2025-01-03 05:00 | 2025-01-30 05:00 | Malfunction | Premature Discharge of Battery; Battery Problem | ACCOLADE SR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THE LONGEVITY ON THIS IMPLANTABLE PACEMAKER HAS DROPPED FROM 11 MONTHS TO LESS THAN 3 MONTHS. DATA WAS SENT FOR ENGINEERING ANALYSIS. ANALYSIS SHOWED THAT THE POWER CONSUMPTION IS INCREASING IN A GRADUAL FASHION. AS OF THIS TIME, THIS ICD REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21266155&pc=LWP | 2124215-2025-04485 | 2025-01-08 05:00 | 2025-01-29 05:00 | Malfunction | Premature Discharge of Battery; Battery Problem | ACCOLADE MRI SR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER WAS SUSPECTED TO HAVE PREMATURE BATTERY DEPLETION (PBD). FURTHERMORE, THERE WERE DIFFERING LONGEVITY ESTIMATES GIVEN OVER THE LAST AND A HALF YEAR FROM 4.5 YEARS TO 1.5 YEARS. DATA ANALYSIS HAS BEEN REQUESTED FROM TECHNICAL SERVICES (TS) AND AT THIS TIME, NO ADDITIONAL INFORMATION HAS BEEN RECEIVED. IT WAS ALSO NOTED ALL LEAD MEASUREMENTS WERE SATISFACTORY AND THE RIGHT VENTRICULAR (RV) LEAD THRESHOLD HAD NOT CHANGED SINCE THE PREVIOUS CHECK AND REMAINED STABLE AT 1V (VOLTS) AT 0.4MS (MILLISECONDS). THE DEVICE REMAINS IN SERVICE AND NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21270325&pc=LWP | 2124215-2025-04658 | 2025-01-02 05:00 | 2025-01-29 05:00 | Malfunction | Premature Discharge of Battery; Battery Problem | ACCOLADE EL DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER EXHIBITED LONGEVITY CALCULATIONS FROM 3.5 YEARS TO 1.5 YEARS REMAINING IN A SPAN OF 9 MONTHS. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS CONFIRMED THE BATTERY APPEARED TO BE DEPLETING MORE QUICKLY THAN EXPECTED. THE POWER CONSUMPTION IS INCREASING IN A GRADUAL FASHION. THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21161768&pc=LWP | 2124215-2025-01548 | 2025-01-05 05:00 | 2025-01-15 05:00 | Malfunction | Premature Discharge of Battery; Battery Problem | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WENT FROM A BATTERY STATUS OF 4 YEARS REMAINING TO A STATUS OF 1 YEAR REMAINING ONE YEAR LATER. THE DEVICE BATTERY WAS SUSPECTED TO BE DEPLETING PREMATURELY. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS CONFIRMED THE BATTERY APPEARED TO BE DEPLETING MORE QUICKLY THAN EXPECTED. DEVICE REPLACEMENT WAS RECOMMENDED, AND A REPLACEMENT INTERVAL WAS DISCUSSED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21164685&pc=LWP | 2124215-2025-01712 | 2025-01-09 05:00 | 2025-01-15 05:00 | Malfunction | Premature Discharge of Battery; Battery Problem | ACCOLADE MRI SR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED OVER THE COURSE OF A YEAR; THIS IMPLANTABLE PACEMAKER LONGEVITY WAS OBSERVED TO HAVE DECREASED FROM 2 YEARS TO LESS THAN 3 MONTHS REMAINING. BOSTON SCIENTIFIC TECHNICAL SERVICES WAS CONTACTED AND RECOMMENDED AN EMERGENT REPLACEMENT PROCEDURE. THE PHYSICIAN SCHEDULED THE PROCEDURE, WHICH HAS NOT YET OCCURRED. AT THIS TIME, THE PACEMAKER REMAINS IMPLANTED AND IN-SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21110626&pc=LWP | 2124215-2025-00438 | 2024-12-12 05:00 | 2025-01-09 05:00 | Malfunction | Premature Discharge of Battery; Battery Problem | ACCOLADE EL DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THE POWER CONSUMPTION ON THIS IMPLANTABLE PACEMAKER HAS INCREASED. DATA WAS SENT FOR ENGINEERING ANALYSIS. ANALYSIS SHOWED THAT THE POWER CONSUMPTION IS INCREASING IN A GRADUAL FASHION. DEVICE REPLACEMENT WAS SUGGESTED. TO DATE, THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21103443&pc=LWP | 2124215-2025-00243 | 2024-12-09 05:00 | 2025-01-08 05:00 | Malfunction | Premature Discharge of Battery; Battery Problem | ACCOLADE EL DR | Tachycardia | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER DEVICE HAS DROPPED IN LONGEVITY FOR 4 YEARS BATTERY REMAINING. DATA ANALYSIS WAS REQUESTED TO HAVE THIS DEVICE ANALYZED WHICH STATED THAT THE POWER CONSUMPTION WAS EXPECTED. THE PATIENT HAS SHORT BURSTS OF ATRIAL TACHYCARDIA AND MIGHT BE CONTINUING TO THE UNEXPECTED DROP. MOREOVER, THE POWER CONSUMPTION WAS INCREASING IN A GRADUAL FASHION WHICH APPEARED TO BE CONSISTENT WITH CAPACITOR DEGRADATION DUE TO HYDROGEN GAS CONTENT IN THE SEALED DEVICE ATMOSPHERE. AS OF THIS TIME, THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21108395&pc=LWP | 2124215-2024-83344 | 2024-12-15 05:00 | 2025-01-08 05:00 | Malfunction | Premature Discharge of Battery; Battery Problem | ACCOLADE MRI SR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITING WAS PREMATURE BATTERY DEPLETION. THE BATTERY HAD DEPLETED BY THREE YEARS AND HALF IN APPROXIMATELY ONE YEAR. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS CONFIRMED THE BATTERY APPEARED TO BE DEPLETING MORE QUICKLY THAN EXPECTED. DEVICE REPLACEMENT WAS RECOMMENDED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. CURRENTLY, EVIDENCE SUGGESTS THAT THIS PRODUCT REMAINS IN-SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21032585&pc=LWP | 2124215-2024-82295 | 2024-12-09 05:00 | 2024-12-29 05:00 | Malfunction | Premature Discharge of Battery; Battery Problem | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS IMPLANTABLE DEVICE WAS SUSPECTED TO BE DEPLETING PREMATURELY. THE DEVICE IS USING 184 PERCENT OF ITS POWER WITH APPROXIMATELY 2.5 YEARS REMAINING LONGEVITY. THE DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: AS NO FURTHER INFORMATION CONCERNING THIS REPORT IS EXPECTED, OUR INVESTIGATION IS COMPLETE. THIS INVESTIGATION WILL BE UPDATED SHOULD FURTHER INFORMATION BE PROVIDED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20956375&pc=LWP | 2124215-2024-79961 | 2024-11-22 05:00 | 2024-12-17 05:00 | Malfunction | Premature Discharge of Battery; Battery Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS PACEMAKER WAS SUSPECTED TO BE DEPLETING PREMATURELY. TECHNICAL SERVICES (TS) ASSESSMENT CONFIRMED THE BATTERY APPEARED TO BE DEPLETING MORE QUICKLY THAN EXPECTED. DEVICE REPLACEMENT WAS RECOMMENDED. THIS DEVICE REMAINED IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20939536&pc=LWP | 2124215-2024-79350 | 2024-12-09 05:00 | 2024-12-16 05:00 | Malfunction | Premature Discharge of Battery; Battery Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THERE WERE CONCERNS OF PREMATURE BATTERY DEPLETION (PBD) ON THIS PACEMAKER DEVICE. THE PATIENT WAS IN ATRIAL FIBRILLATION (AF). THE HEALTH CARE PROFESSIONAL (HCP) HAD BEEN CONCERNED FOR THE BATTERY LONGEVITY. THE PATIENT WILL BE SEEN IN NEAR FUTURE TO CHECK FOR REPLACEMENT PROCEDURE. THIS DEVICE CURRENTLY REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20940142&pc=LWP | 2124215-2024-79416 | 2024-12-09 05:00 | 2024-12-16 05:00 | Malfunction | Premature Discharge of Battery; Battery Problem | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SHOWED A BATTERY STATUS OF 5.5 YEARS REMAINING. THE DEVICE BATTERY WAS SUSPECTED TO BE DEPLETING PREMATURELY. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS CONFIRMED THE BATTERY APPEARED TO BE DEPLETING MORE QUICKLY THAN EXPECTED. DEVICE REPLACEMENT WAS RECOMMENDED, AND A REPLACEMENT INTERVAL WAS DISCUSSED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20925323&pc=LWP | 2124215-2024-78998 | 2024-12-03 05:00 | 2024-12-13 05:00 | Malfunction | Premature Discharge of Battery; Battery Problem | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE BATTERY OF THIS PACEMAKER WAS SUSPECTED TO BE DEPLETING PREMATURELY. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. DATA ANALYSIS CONFIRMED THE BATTERY APPEARED TO BE DEPLETING MORE QUICKLY THAN EXPECTED. A DEVICE REPLACEMENT INTERVAL WAS PROVIDED. THIS DEVICE REMAINED IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20912821&pc=LWP | 2124215-2024-78617 | 2024-12-05 05:00 | 2024-12-12 05:00 | Malfunction | Premature Discharge of Battery; Battery Problem | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER DEVICE IS SUSPECTED OF BEHAVING IN A MANNER CONSISTENT WITH PREMATURE BATTERY DEPLETION (PBD). IN 2022, THE REMAINING BATTERY LONGEVITY WAS 12.5 YEARS, AND CURRENTLY THE REMAINING BATTERY LONGEVITY IS NOW 2 YEARS. A REQUEST WAS MADE TO HAVE DATA FROM THIS DEVICE ANALYZED. A TECHNICAL SERVICE (TS) ADVISED TO SEND THEM THE DEVICE DATA FOR REVIEW. AT THIS TIME, NO INFORMATION HAS BEEN RECEIVED TO REVIEW. THE DEVICE REMAINS IMPLANTED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. AT THIS TIME EVIDENCE SUGGESTS THAT THIS DEVICE REMAINS IMPLANTED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20914924&pc=LWP | 2124215-2024-78722 | 2024-11-19 05:00 | 2024-12-12 05:00 | Malfunction | Premature Discharge of Battery; Battery Problem | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED A SIGNIFICANT DECREASE IN THE REMAINING LONGEVITY, DOWN FROM 10 YEARS IN 2021 TO NINE MONTHS IN 2024. PREMATURE BATTERY DEPLETION WAS SUSPECTED. DEVICE DATA WAS SUBMITTED TO TECHNICAL SERVICES FOR REVIEW. ANALYSIS OF THE DATA CONFIRMED THE DEVICE WAS DEPLETING FASTER THAN EXPECTED, AND DEVICE REPLACEMENT WAS RECOMMENDED FOR WITHIN 90 DAYS. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THE DEVICE REMAINS IMPLANTED AND IN SERVICE. Manufacturer Narrative: THIS REPORT WILL BE UPDATED WHEN ADDITIONAL INFORMATION IS RECEIVED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21225948&pc=LWP | 2124215-2025-03339 | 2024-12-30 05:00 | 2025-01-23 05:00 | Malfunction | Premature Discharge of Battery; Incorrect, Inadequate or Imprecise Result or Readings | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT DURING PRE-IMPLANT TESTING, INTERROGATION OF THIS PACEMAKER FOUND THE DEVICE TO BE IN SAFETY MODE. PREMATUER BATTERY DEPLETION WAS SUSPECTED. THE DEVICE WAS NOT USED; THE IMPLANT PROCEDURE WAS CANCELLED BEFORE THE PATIENT WAS SEDATED AND WILL BE RESCHEDULED AT A LATER DATE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THE DEVICE WAS EXPECTED TO BE RETURNED FOR ANALYSIS. Manufacturer Narrative: THE PRODUCT IS EXPECTED TO BE RETURNED FOR ANALYSIS. THIS REPORT WILL BE UPDATED UPON RETURN AND COMPLETION OF ANALYSIS.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21154788&pc=LWP | 2124215-2024-84032 | 2024-12-19 05:00 | 2025-01-14 05:00 | Malfunction | Premature Discharge of Battery; Over-Sensing | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS SUSPECTED TO BE EXHIBITING PREMATURE BATTERY DEPLETION. TECHNICAL SERVICES REVIEWED THE DEVICE DATA AND CONFIRMED DEVICE WAS DEPLETING NORMALLY WITH STABLE POWER CONSUMPTION AND WAS ON TRACK TO MEET EXPECTED LONGEVITY ESTIMATES BASED ON CURRENT SETTINGS. IT WAS DISCUSSED THAT THE DEVICE HAD STORED MANY EPISODES INDICATING HIGH-RATE ATRIAL SENSING HAD OCCURRED, WHICH CAN CAUSE AN INCREASE IN POWER CONSUMPTION. DURING THE REVIEW, TECHNICAL SERVICES NOTED FAR FIELD OVERSENSING OF VENTRICULAR PACED BEATS ON THE ATRIAL CHANNEL. THIS WAS UNRELATED TO THE BATTERY CONCERNS. TECHNICAL SERVICES RECOMMENDED MEASURING THE SIGNALS AND ADJUSTING THE ATRIAL SENSITIVITY TO AVOID OVERSENSING THE VENTRICULAR PACED BEATS. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THE DEVICE REMAINS IMPLANTED AND IN SERVICE.Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS SUSPECTED TO BE EXHIBITING PREMATURE BATTERY DEPLETION. TECHNICAL SERVICES REVIEWED THE DEVICE DATA AND CONFIRMED DEVICE WAS DEPLETING NORMALLY WITH STABLE POWER CONSUMPTION AND WAS ON TRACK TO MEET EXPECTED LONGEVITY ESTIMATES BASED ON CURRENT SETTINGS. IT WAS DISCUSSED THAT THE DEVICE HAD STORED MANY EPISODES INDICATING HIGH RATE ATRIAL SENSING HAD OCCURRED, WHICH CAN CAUSE AN INCREASE IN POWER CONSUMPTION. DURING THE REVIEW, TECHNICAL SERVICES NOTED FAR FIELD OVERSENSING OF VENTRICULAR PACED BEATS ON THE ATRIAL CHANNEL. THIS WAS UNRELATED TO THE BATTERY CONCERNS. TECHNICAL SERVICES RECOMMENDED MEASURING THE SIGNALS AND ADJUSTING THE ATRIAL SENSITIVITY TO AVOID OVERSENSING THE VENTRICULAR PACED BEATS. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THE DEVICE REMAINS IMPLANTED AND IN SERVICE. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THIS PRODUCT HAS NOT BEEN RETURNED TO BOSTON SCIENTIFIC, AND AS A RESULT, LABORATORY ANALYSIS COULD NOT BE CONDUCTED. THE ASSOCIATED INVESTIGATION DETERMINED THAT THIS DEVICE EXHIBITED SENSING OF CHRONIC HIGH ATRIAL RATES WHICH MAY IMPACT BATTERY LONGEVITY OVER TIME. THE OBSERVATION OF FAR FIELD OVERSENSING OF VENTRICULAR PACED BEATS ON THE ATRIAL CHANNEL IS A KNOWN INHERENT RISK WITH THE USE OF THIS PRODUCT. PLEASE REFER TO THE DESCRIPTION FOR MORE INFORMATION REGARDING THE SPECIFIC CIRCUMSTANCES OF THIS EVENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21175293&pc=LWP | 2124215-2025-01964 | 2024-12-30 05:00 | 2025-01-16 05:00 | Malfunction | Premature Discharge of Battery; Telemetry Discrepancy | ACCOLADE DR | Chest Pain; Discomfort | N970003 | 1 | Event Description: IT WAS REPORTED BY THE PATIENT THAT THIS PACEMAKER HAD ISSUES TO BE INTERROGATED. IN ADDITION, THERE WERE CONCERNS OF PREMATURE BATTERY DEPLETION (PBD) OR BATTERY ISSUES. THE LAST DEVICE INTERROGATION WAS SEVERAL YEARS PRIOR. THE PATIENT FELT CHEST PAIN AND DISCOMFORT. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.Event Description: IT WAS REPORTED BY THE PATIENT THAT THIS PACEMAKER HAD ISSUES TO BE INTERROGATED. IN ADDITION, THERE WERE CONCERNS OF PREMATURE BATTERY DEPLETION (PBD) OR BATTERY ISSUES. THE LAST DEVICE INTERROGATION WAS SEVERAL YEARS PRIOR. THE PATIENT FELT CHEST PAIN AND DISCOMFORT. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE. ADDITIONAL INFORMATION WAS RECEIVED, THE CLINICAL REPRESENTATIVE INDICATED THAT THERE WERE NO SIGNS OF PBD. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21105772&pc=LWP | 2124215-2025-00290 | 2024-09-18 04:00 | 2025-01-08 05:00 | Malfunction | Signal Artifact/Noise; Failure to Capture; High impedance; Over-Sensing; Pacing Problem; Data Problem | ACCOLADE MRI DR | Unspecified Heart Problem | P150012 | 1 | Event Description: IT WAS REPORTED THAT DURING IMPLANT PROCEDURE, THIS PACEMAKER DEVICE WAS OBSERVED TO HAVE RECEIVED AN ALERT THAT INDICATED HIGH OUT OF RANGE IMPEDANCE MEASUREMENTS ON THE RIGHT ATRIAL (RA) AND RIGHT VENTRICULAR (RV) LEADS. IT WAS NOTED THAT THE LEADS WERE NON-BOSTON SCIENTIFIC PRODUCTS. THE DAILY IMPEDANCE MEASUREMENTS WERE REVIEWED AND FOUND TO HAVE MEASUREMENTS FROM ABOUT 10 DAYS PRIOR TO THE DEVICE BEING IMPLANTED. IT WAS DETERMINED THAT THE DEVICE HAD BEEN OUT OF STORAGE MODE PRIOR TO IMPLANT WHICH LED TO THE FALSE HIGH OUT OF RANGE IMPEDANCE MEASUREMENTS RECORDED. THE PRESENTING ELECTROGRAM (EGM) SHOWS APPROPRIATE RIGHT ATRIAL (RA) LAND RIGHT VENTRICULAR (RV) SENSING AND VENTRICULAR CAPTURE WAS CONFIRMED. AT THIS TIME, THE PACEMAKER REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. SUBSEQUENT ADDITIONAL INFORMATION WAS RECEIVED THAT REPORTED DURING A FOLLOW UP CLINIC VISIT, THE HEALTH CARE PROFESSIONAL (HCP) CALLED TECHNICAL SERVICES (TS) FOR FURTHER REVIEW OF THE PACEMAKER PRESENTING ELECTROGRAM (EGM) DUE TO CONCERNS OF OVERSENSING NOISE, PACING INHIBITION AND POSSIBLE LOSS OF CAPTURE (LOC) ON THE NON-BOSTON SCIENTIFIC RIGHT ATRIAL (RA) AND RIGHT VENTRICULAR (RV) LEADS. THE HCP NOTED THAT THE PATIENT HAD REPORTED EXPERIENCING PALPITATION SYMPTOMS. UPON REVIEW, THE PRESENTING EGM SHOWED THE NON-BOSTON SCIENTIFIC RV LEAD EXHIBITED OVERSENSING NOISE WITH PACING INHIBITION WITH A SUDDEN INCREASE IN PACING IMPEDANCES FROM 429 TO 971 OHMS. TS NOTED POSSIBLE LOSS OF CAPTURE (LOC) ON THE NON-BOSTON SCIENTIFIC RA LEAD BUT COULD NOT SEE ANY OVERSENSING ON THE EVENTS REVIEWED. TS DISCUSSED WITH THE HCP PROGRAMMING OPTIMIZATION OPTIONS. AT THIS TIME, THE PACEMAKER AND NON-BOSTON SCIENTIFIC RA AND RV LEADS REMAIN IN SERVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE WAS NOT RETURNED FOR ANALYSIS. PI ANALYSIS WAS COMPLETED ON THE DEVICE USING AVAILABLE INFORMATION. THE ALLEGATION(S) IN THIS COMPLAINT HAVE NOT BEEN CONFIRMED AS THE PRODUCT HAS NOT BEEN RETURNED FOR ANALYSIS. NO DEVICE ISSUES ARE INDICATED BASED ON AVAILABLE INFORMATION.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20908783&pc=LWP | 2124215-2024-78489 | 2024-12-03 05:00 | 2024-12-11 05:00 | Malfunction | Signal Artifact/Noise; Failure to Capture; High impedance; Over-Sensing; Use of Device Problem; High Capture Threshold | ACCOLADE MRI DR | Presyncope; Asystole | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM EXHIBITED HIGH OUT OF RANGE PACE IMPEDANCE MEASUREMENTS ON THE RIGHT VENTRICULAR (RV) LEAD RESULTING IN A LEAD SAFETY SWITCH. A LEAD REVISION WAS SCHEDULED; HOWEVER, IT DID NOT OCCUR DUE TO UNKNOWN REASONS. ADDITIONALLY, OVERSENSING OF NOISY SIGNALS WAS OBSERVED RESULTING IN PACING INHIBITION OF 4.5 SECONDS. TS DISCUSSED THIS WAS LIKELY DUE TO MYOPOTENTIALS. ADDITIONALLY, THE RV LEAD WAS SUSPECTED TO EXHIBIT LOSS OF CAPTURE (LOC). BOSTON SCIENTIFIC TECHNICAL SERVICES (TS) WAS CONSULTED AND UPON REVIEW, HIGH CAPTURE THRESHOLDS WERE NOTED. THE RIGHT VENTRICULAR AUTOMATIC THRESHOLD (RVAT) TEST WAS FOUND TO BE UNSUCCESSFUL DUE TO LOW ER. THE PATIENT WAS REPORTED TO BE NEAR SYNCOPAL AT TIMES WHEN PERFORMING PHYSICAL ACTIVITY. NO ADVERSE PATIENT EFFECTS WERE REPORTED. AT THIS TIME, THIS DEVICE SYSTEM REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21199762&pc=LWP | 2124215-2025-02668 | 2024-12-30 05:00 | 2025-01-21 05:00 | Malfunction | Signal Artifact/Noise; Failure to Capture; Over-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER EXHIBITED FARFIELD OVERSENSING OF RIGHT VENTRICULAR SENSE (RVS) ON THE ATRIAL CHANNEL. TECHNICAL SERVICES RECOMMENDED REPROGRAMMING DEVICE TO FIXED SENSITIVITY AND EXTENDING ATRIAL BLANK OF RVS. ADDITIONALLY, THERE WAS INTERMITTENT LOSS OF CAPTURE ON THE RV CHANNEL. THE DEVICE CURRENTLY REMAINS IN SERVICE AND NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21223768&pc=LWP | 2124215-2025-03220 | 2024-09-01 04:00 | 2025-01-23 05:00 | Malfunction | Signal Artifact/Noise; Failure to Capture; Over-Sensing; Low impedance; Device Sensing Problem; Impedance Problem | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THE ATRIAL INTRINSIC AMPLITUDES WERE OUT OF RANGE AND ATRIAL LEAD NOISE/OVERSENSING WAS NOTED IN THE ATRIAL TACHY RESPONSE (ATR) EPISODES. IN ADDITION THERE WERE NOTED SPIKES IN THE VENTRICULAR PACE IMPEDANCES WITH VALUES BEING WITHIN RANGE. IT WAS ALSO NOTED THAT THE AUTOMATIC SAFETY SWITCH WAS TRIGGERED FROM BIPOLAR TO UNIPOLAR CONFIGURATION WHEN IT COMES TO THIS DEVICE DUE TO LOW OUT OF RANGE PACE IMPEDANCE MEASUREMENTS ON THE RIGHT ATRIAL CHANNEL. TECHNICAL SERVICES (TS) REVIEW OF THE DEVICE DATA TOOK PLACE AND TS RECOMMENDED TO ACCESS THE COMPETITOR ATRIAL AND VENTRICULAR LEAD INTEGRITY. ADDITIONAL INFORMATION NOTED THAT MOST RECENTLY THE ATRIAL AUTOMATIC THRESHOLDS WAS DETECTED AS SUSPENDED AND PRESENTING ELECTROGRAM (EGM) SHOWED INAPPROPRIATE ATRIAL SENSING FOLLOWING THE RECENT LEAD POLARITY SWITCH. IT WAS NOT POSSIBLE TO CONFIRM CAPTURE ON THE PRESENTING EGM WITH NO SUCCESSFULLY THRESHOLD MEASUREMENTS OBTAINED IN JANUARY 2025. IT WAS DISCUSSED TO HAVE THE PATIENT BE SEEN IN CLINIC AND CHANGE THE ATRIAL SENSING BACK TO BIPOLAR CONFIGURATION. AT THIS TIME THE DEVICE REMAINS IMPLANTED. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20957134&pc=LWP | 2124215-2024-80026 | 2024-11-18 05:00 | 2024-12-17 05:00 | Malfunction | Signal Artifact/Noise; Failure to Capture; Over-Sensing; Pacing Problem; Connection Problem | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM EXHIBITED BRADY PACING NOT DELIVERED WHEN REQUIRED FROM THE RIGHT VENTRICULAR (RV) LEAD AND NOISE AND OVERSENSING. LOSS OF CAPTURE WAS NOTED. THE ISSUE WAS RESOLVED WITH REPROGRAMMING. THIS PACEMAKER SYSTEM REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM EXHIBITED BRADY PACING NOT DELIVERED WHEN REQUIRED FROM THE RIGHT VENTRICULAR (RV) LEAD AND NOISE AND OVERSENSING. NOISE WAS DUE TO A SUSPECTED DISCONNECTION. LOSS OF CAPTURE WAS NOTED. THE ISSUE WAS RESOLVED WITH REPROGRAMMING. THIS PACEMAKER SYSTEM REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21284151&pc=LWP | 2124215-2025-05047 | 2025-01-17 05:00 | 2025-01-31 05:00 | Malfunction | Signal Artifact/Noise; Failure to Capture; Pacing Problem | ACCOLADE EL DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THE HEALTH CARE PROFESSIONAL (HCP) CALLED FOR REVIEW OF DEVICE DATA FOR THIS PATIENT WITH A PACEMAKER. TECHNICAL SERVICES (TS) REVIEWED AND FOUND RIGHT VENTRICULAR (RV) LOSS OF CAPTURE. THE NEXT BEAT IS DECLARED FUSION HOWEVER, THERE WAS NO DEFLECTION ON THE ELECTROGRAM (EGM). TS WAS NOT SURE IF IT WAS TRULY FUSION OR NOT. TS REVIEWED LEAD TRENDS, AND ALL SEEMED TO BE STABLE. ADDITIONALLY, THERE WAS A STORED PACEMAKER MEDIATED TACHYCARDIA (PMT) EPISODE AND NOISE WAS FOUND ON THE RV EGM. THERE WAS NO PACING INHIBITION. THE PATIENT IS SCHEDULED FOR A GENERATOR CHANGE OUT AND TS RECOMMENDED TO EVALUATE THE RV LEAD. AT THIS TIME, THE DEVICE AND NON- BOSTON SCIENTIFIC RV LEAD REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. IT WAS REPORTED THAT THE HEALTH CARE PROFESSIONAL (HCP) CALLED FOR REVIEW OF DEVICE DATA FOR THIS PATIENT WITH A PACEMAKER. TECHNICAL SERVICES REVIEWED AND FOUND THERE WAS RIGHT VENTRICULAR (RV) LOSS OF CAPTURE HOWEVER, THE NEXT BEAT IS DECLARED FUSION BUT THERE WAS NO DEFLECTION ON THE ELECTROGRAM (EGM). TS INFORMED IT WAS HARD TO TELL IF IT WAS TRULY FUSION OR NOT. RV LEAD TRENDS APPEARED TO BE STABLE. TS DID FIND A STORED PACEMAKER MEDIATED TACHYCARDIA (PMT) EPISODE AND DID NOTE THERE WAS NOISE ON THE RV EGM. THERE WAS NO PACING INHIBITION. TS RECOMMENDED EVALUATING THE RV LEAD AT THE TIME OF THE SCHEDULED GENERATOR CHANGE. AT THIS TIME, THE DEVICE AND NON-BOSTON SCIENTIFIC RV LEAD REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21248486&pc=LWP | 2124215-2025-03913 | 2025-01-13 05:00 | 2025-01-27 05:00 | Malfunction | Signal Artifact/Noise; High impedance | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXHIBITING HIGH PACING IMPEDANCE OUT OF RANGE AND NOISY SIGNALS. THE LEAD REMAINS IN SERVICE. NO INFORMATION IS AVAILABLE AT THE MOMENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21275174&pc=LWP | 2124215-2025-04771 | 2023-06-23 04:00 | 2025-01-30 05:00 | Malfunction | Signal Artifact/Noise; High impedance; Over-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED HIGH PACING IMPEDANCE OUT OF RANGE (OOR) MEASUREMENTS ON THE RIGHT ATRIAL (RA) LEAD, WHICH TRIGGERED TO LEAD SAFETY SWITCH (LSS) FROM BIPOLAR TO UNIPOLAR MODE. THE PATIENT WAS BROUGHT TO AN OFFICE CHECK AND THE LEAD WAS REPROGRAMMED TO UNIPOLAR MODE AND THE IMPEDANCE MEASUREMENTS WERE IN RANGE, THEN REPROGRAMMED TO BIPOLAR MODE AND THE MEASUREMENTS WERE STILL IN RANGE. NOISE WAS ALSO OBSERVED DURING ISOMETRIC TESTING. THE ELECTROCARDIOGRAM SHOWED A LOT OF ATRIAL NOISE BEFORE, AND AFTER THE LSS. ADDITIONALLY, IT WAS NOTED THAT THE PACEMAKER RECORDED SIGNAL ARTIFACT MONITORING (SAM) EPISODES DUE TO SUSPECTED NOISY SIGNALS FROM ELECTROMAGNETIC INTERFERENCE (EMI). THIS RESULTED IN THE MINUTE VENTILATION (MV) FEATURE BEING AUTOMATICALLY DISABLED. THE RA LEAD IS 20 YEARS OLD. IT WAS SUSPECTED THAT THE RA LEAD IS LIKELY TO HAVE LEAD INTEGRITY ISSUES CAUSING NOISE AND HIGH OOR IMPEDANCE INTERMITTENTLY IN BIPOLAR. CURRENTLY, THIS PRODUCT REMAINS IN SERVICE AND NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21279770&pc=LWP | 2124215-2025-04909 | 2025-01-09 05:00 | 2025-01-30 05:00 | Malfunction | Signal Artifact/Noise; High impedance; Over-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: IT WAS REPORTED THAT THE RIGHT VENTRICULAR (RV) LEAD EXHIBITED A HIGH OUT OF RANGE PACING IMPEDANCE MEASUREMENT GREATER THAN 3000 OHMS TRIGGERING A LEAD SAFETY SWITCH (LSS) TO UNIPOLAR. IT WAS ALSO NOTED THE RV LEAD HAS EXHIBITED NOISE ISSUES SINCE IMPLANT THREE MONTHS AGO. THE PATIENT WAS SEEN IN-CLINIC AND THE BIPOLAR IMPEDANCE AND THRESHOLD MEASUREMENTS WERE GOOD. TECHNICAL SERVICES (TS) WAS CONSULTED AND ADVISED IN NOVEMBER OF LAST YEAR THIS IMPLANTABLE PACEMAKER RECORDED A SIGNAL ARTIFACT MONITORING (SAM) EPISODE DUE TO OVERSENSING OF NOISE ON THE RV LEAD RESULTING IN THE MINUTE VENTILATION (MV) FEATURE BEING AUTOMATICALLY DISABLED. TS PROVIDED TROUBLESHOOTING AND DEVICE OPTIMIZATION RECOMMENDATIONS. AT THIS TIME, THE LEAD AND DEVICE REMAIN IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21126921&pc=LWP | 2124215-2025-00833 | 2023-12-09 05:00 | 2025-01-10 05:00 | Malfunction | Signal Artifact/Noise; High impedance; Over-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT DURING A PROCEDURE THIS PACEMAKER INAPPROPRIATELY STORED FOUR SIGNAL ARTIFACT MONITOR (SAM) EPISODES AND DISABLED THE MINUTE VENTILATION (MV). THE EVENTS RECORDED HIGH OUT OF RANGE PACING IMPEDANCE MEASUREMENTS AND NOISE OVERSENSING ON THE RIGHT ATRIAL (RA) AND RIGHT VENTRICULAR (RV) LEADS. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21106395&pc=LWP | 2124215-2024-83230 | 2024-12-13 05:00 | 2025-01-08 05:00 | Malfunction | Signal Artifact/Noise; High impedance; Over-Sensing | ACCOLADE EL DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM EXHIBITED HIGH OUT OF RANGE PACING IMPEDANCES MEASURING 2195 OHMS, WHICH LED TO A LEAD SAFETY SWITCH (LSS). TECHNICAL SERVICES (TS) WAS CONSULTED AND UPON REVIEW INAPPROPRIATE ATRIAL TACHY RESPONSE (ATR) EPISODES CAUSED BY NOISE OVERSENSING WERE FOUND. TS ADVISED ON FURTHER IN OFFICE REVIEW OF THE SYSTEM. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS SYSTEM REMAINS IN SERVICE.Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM EXHIBITED HIGH OUT OF RANGE PACING IMPEDANCES MEASURING 2195 OHMS, WHICH LED TO A LEAD SAFETY SWITCH (LSS). TECHNICAL SERVICES (TS) WAS CONSULTED AND UPON REVIEW INAPPROPRIATE ATRIAL TACHY RESPONSE (ATR) EPISODES CAUSED BY NOISE OVERSENSING WERE FOUND. TS ADVISED ON FURTHER IN OFFICE REVIEW OF THE SYSTEM. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS SYSTEM REMAINS IN SERVICE. ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT THIS DEVICE WAS EXPLANTED AND REPLACED PROPHYLACTICALLY. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: THIS PRODUCT HAS NOT BEEN RETURNED TO BOSTON SCIENTIFIC, AND AS A RESULT, LABORATORY ANALYSIS COULD NOT BE CONDUCTED. THE ASSOCIATED INVESTIGATION DETERMINED THAT THIS DEVICE EXHIBITED INTERMITTENT CHANGES IN IMPEDANCE MEASUREMENTS WITH NO CONCLUSIVE EVIDENCE OF A PRODUCT PERFORMANCE ISSUE OR INADEQUATE LEAD-TO-DEVICE CONNECTION; PLEASE REFER TO THE DESCRIPTION FOR MORE INFORMATION REGARDING THE SPECIFIC CIRCUMSTANCES OF THIS EVENT. ENGINEERING ANALYSIS AND TESTING OF RETURNED PRODUCTS HAS DETERMINED THAT REPEATED, SMALL MOVEMENTS OF THE LEAD TERMINAL RING ALONG WITH VARIANCES IN THE CONNECTION CONTACT FORCE AND SURFACE ROUGHNESS CAN IMPACT THE CONNECTION BETWEEN THE SPRING CONTACT AND THE LEAD RING, RESULTING IN INTERMITTENT CHANGES IN IMPEDANCE MEASUREMENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20945883&pc=LWP | 2124215-2024-79529 | 2024-12-05 05:00 | 2024-12-16 05:00 | Malfunction | Signal Artifact/Noise; High impedance; Over-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED A LEAD SAFETY SWITCH (LSS) ON THE RIGHT ATRIAL (RA) LEAD AND THE HEALTH CARE PROFESSIONAL (HCP) REQUESTED A FURTHER REVIEW OF THE DATA. TECHNICAL SERVICES (TS) WAS CONSULTED, AND IT WAS FOUND THAT THIS PACEMAKER RECORDED A SIGNAL ARTIFACT MONITOR (SAM) EPISODE DUE TO OVERSENSING ARTIFACT RELATED TO THE MINUTE VENTILATION (MV) FEATURE SIGNAL ON THE RA CHANNEL. IN ADDITION, THE LSS OCCURRED DUE TO HIGH OUT OF RANGE PACING IMPEDANCE MEASUREMENTS ON THE RA LEAD. FURTHER TESTING WAS RECOMMENDED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED A LEAD SAFETY SWITCH (LSS) ON THE RIGHT ATRIAL (RA) LEAD AND THE HEALTH CARE PROFESSIONAL (HCP) REQUESTED A FURTHER REVIEW OF THE DATA. TECHNICAL SERVICES (TS) WAS CONSULTED, AND IT WAS FOUND THAT THIS PACEMAKER RECORDED A SIGNAL ARTIFACT MONITOR (SAM) EPISODE DUE TO OVERSENSING ARTIFACT RELATED TO THE MINUTE VENTILATION (MV) FEATURE SIGNAL ON THE RA CHANNEL. IN ADDITION, THE LSS OCCURRED DUE TO HIGH OUT OF RANGE PACING IMPEDANCE MEASUREMENTS ON THE RA LEAD. FURTHER TESTING WAS RECOMMENDED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE. Manufacturer Narrative: THIS PRODUCT HAS NOT BEEN RETURNED TO BOSTON SCIENTIFIC, AND AS A RESULT, LABORATORY ANALYSIS COULD NOT BE CONDUCTED. THE ASSOCIATED INVESTIGATION DETERMINED THAT THIS DEVICE EXHIBITED INTERMITTENT CHANGES IN IMPEDANCE MEASUREMENTS WITH NO CONCLUSIVE EVIDENCE OF A PRODUCT PERFORMANCE ISSUE OR INADEQUATE LEAD-TO-DEVICE CONNECTION; PLEASE REFER TO THE DESCRIPTION FOR MORE INFORMATION REGARDING THE SPECIFIC CIRCUMSTANCES OF THIS EVENT. ENGINEERING ANALYSIS AND TESTING OF RETURNED PRODUCTS HAS DETERMINED THAT REPEATED, SMALL MOVEMENTS OF THE LEAD TERMINAL RING ALONG WITH VARIANCES IN THE CONNECTION CONTACT FORCE AND SURFACE ROUGHNESS CAN IMPACT THE CONNECTION BETWEEN THE SPRING CONTACT AND THE LEAD RING, RESULTING IN INTERMITTENT CHANGES IN IMPEDANCE MEASUREMENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21146169&pc=LWP | 2124215-2025-01165 | 2024-12-12 05:00 | 2025-01-13 05:00 | Malfunction | Signal Artifact/Noise; High impedance; Over-Sensing; Battery Problem | ACCOLADE DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT THERE WERE CONCERNS OF THIS PACEMAKER WAS DEPLETING FASTER THAN EXPECTED AS THE DEVICE SHOWED MORE THAN THREE MONTHS OF BATTERY REMAINING IN DECEMBER AND TWO WEEKS LATER ELECTIVE REPLACEMENT INDICATOR (ERI) WAS REACHED. IN ADDITION, A SUDDEN HIGH WITHIN RANGE IMPEDANCE MEASUREMENT WAS NOTED ON THE RIGHT ATRIAL (RA) CHANNEL. TECHNICAL SERVICES (TS) WAS CONSULTED, AND IT WAS DETERMINED THAT THE ERI WAS APPROPRIATE AS IT APPEARED IN AUGUST. A MINUTE VENTILATION (MV) SIGNAL WAS ALSO NOTED, WITH INAPPROPRIATE ATRIAL TACHY RESPONSE (ATR) EPISODES STORED DUE TO NOISE OVERSENSING. FURTHER MONITORING WAS RECOMMENDED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.Event Description: IT WAS REPORTED THAT THERE WERE CONCERNS OF THIS PACEMAKER WAS DEPLETING FASTER THAN EXPECTED AS THE DEVICE SHOWED MORE THAN THREE MONTHS OF BATTERY REMAINING IN DECEMBER AND TWO WEEKS LATER ELECTIVE REPLACEMENT INDICATOR (ERI) WAS REACHED. IN ADDITION, A SUDDEN HIGH WITHIN RANGE IMPEDANCE MEASUREMENT WAS NOTED ON THE RIGHT ATRIAL (RA) CHANNEL. TECHNICAL SERVICES (TS) WAS CONSULTED, AND IT WAS DETERMINED THAT THE ERI WAS APPROPRIATE AS IT APPEARED IN AUGUST. A MINUTE VENTILATION (MV) SIGNAL WAS ALSO NOTED, WITH INAPPROPRIATE ATRIAL TACHY RESPONSE (ATR) EPISODES STORED DUE TO NOISE OVERSENSING. FURTHER MONITORING WAS RECOMMENDED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE. Manufacturer Narrative: THIS PRODUCT HAS NOT BEEN RETURNED TO BOSTON SCIENTIFIC, AND AS A RESULT, LABORATORY ANALYSIS COULD NOT BE CONDUCTED. THE ASSOCIATED INVESTIGATION DETERMINED THAT THIS DEVICE EXHIBITED INTERMITTENT CHANGES IN IMPEDANCE MEASUREMENTS WITH NO CONCLUSIVE EVIDENCE OF A PRODUCT PERFORMANCE ISSUE OR INADEQUATE LEAD-TO-DEVICE CONNECTION; PLEASE REFER TO THE DESCRIPTION FOR MORE INFORMATION REGARDING THE SPECIFIC CIRCUMSTANCES OF THIS EVENT. ENGINEERING ANALYSIS AND TESTING OF RETURNED PRODUCTS HAS DETERMINED THAT REPEATED, SMALL MOVEMENTS OF THE LEAD TERMINAL RING ALONG WITH VARIANCES IN THE CONNECTION CONTACT FORCE AND SURFACE ROUGHNESS CAN IMPACT THE CONNECTION BETWEEN THE SPRING CONTACT AND THE LEAD RING, RESULTING IN INTERMITTENT CHANGES IN IMPEDANCE MEASUREMENT.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21114394&pc=LWP | 2124215-2024-83441 | 2024-10-12 04:00 | 2025-01-09 05:00 | Malfunction | Signal Artifact/Noise; High impedance; Over-Sensing; High Capture Threshold | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER DEVICE AND ASSOCIATED RIGHT VENTRICULAR (RV) LEAD EXHIBITED HIGH OUT OF RANGE PACING IMPEDANCE MEASUREMENTS GREATER THAN 3000 OHMS, HIGH PACING THRESHOLD VALUES AND NOISE OVERSENSING. IT IS BELIEVED THAT THESE OBSERVATIONS COULD BE RELATED TO AN RV LEAD INTEGRITY CONCERN. BOSTON SCIENTIFIC TECHNICAL SERVICES (TS) RECOMMENDED LEAD EVALUATION OR A REVISION PROCEDURE, BUT AT THIS TIME, THERE IS NO EVIDENCE TO SUGGEST THAT THIS INTERVENTION HAS BEEN PERFORMED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THE DEVICE REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21160473&pc=LWP | 2124215-2025-01510 | 2024-12-22 05:00 | 2025-01-15 05:00 | Malfunction | Signal Artifact/Noise; High impedance; Over-Sensing; Pacing Problem | ACCOLADE MRI DR | Fall | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS PACEMAKER SYSTEM VISITED THE EMERGENCY ROOM (ER) AFTER A FALL. THE HEALTH CARE PROFESSIONAL (HCP) PLACED A CALL TO TECHNICAL SERVICES (TS) FOR FURTHER REVIEW OF THE PACEMAKER PRESENTING ELECTROGRAM (EGM). UPON REVIEW, OVERSENSING NOISE AND PACING INHIBITION ON THE RIGHT VENTRICULAR (RV) LEAD WAS OBSERVED. THE PRESENTING EGM ALSO SHOWED A STORED SIGNAL ARTIFACT MONITOR (SAM) EPISODE DUE TO HIGH OUT OF RANGE PACING IMPEDANCES ON THE RV LEAD SENSORS WERE DETECTED AT THE TIME OF ARTIFACT DETECTION WHICH LED TO THE SAM FEATURE TO DISABLE THE MINUTE VENTILATION (MV) SENSOR. TS DISCUSSED REVIEW FINDINGS WITH THE HCP AND RECOMMENDED FOR THE DEVICE TREATING PHYSICIAN TO BE CONSULTED FOR FURTHER EVALUATION. AT THIS TIME, THE PACEMAKER AND RV LEAD REMAIN IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS PACEMAKER SYSTEM VISITED THE EMERGENCY ROOM (ER) AFTER A FALL. THE HEALTH CARE PROFESSIONAL (HCP) PLACED A CALL TO TECHNICAL SERVICES (TS) FOR FURTHER REVIEW OF THE PACEMAKER PRESENTING ELECTROGRAM (EGM). UPON REVIEW, OVERSENSING NOISE AND PACING INHIBITION ON THE RIGHT VENTRICULAR (RV) LEAD WAS OBSERVED. THE PRESENTING EGM ALSO SHOWED A STORED SIGNAL ARTIFACT MONITOR (SAM) EPISODE DUE TO HIGH OUT OF RANGE PACING IMPEDANCES ON THE RV LEAD SENSORS WERE DETECTED AT THE TIME OF ARTIFACT DETECTION WHICH LED TO THE SAM FEATURE TO DISABLE THE MINUTE VENTILATION (MV) SENSOR. TS DISCUSSED REVIEW FINDINGS WITH THE HCP AND RECOMMENDED FOR THE DEVICE TREATING PHYSICIAN TO BE CONSULTED FOR FURTHER EVALUATION. AT THIS TIME, THE PACEMAKER AND RV LEAD REMAIN IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. SUBSEQUENT ADDITIONAL INFORMATION WAS RECEIVED THAT REPORTED THAT THE PATIENT WITH THIS PACEMAKER SYSTEM UNDERWENT A PROCEDURE WHERE THIS PACEMAKER WAS BRADY THERAPY PROGRAMMING WAS TURNED OFF AND A NEW DEVICE WAS IMPLANTED AS A REPLACEMENT. ADDITIONALLY, DURING A FOLLOW UP VISIT, NOISE SIGNALS WERE OBSERVED. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THE PACEMAKER AND RV LEAD REMAIN IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21164037&pc=LWP | 2124215-2025-01642 | 2024-10-22 04:00 | 2025-01-15 05:00 | Malfunction | Signal Artifact/Noise; High impedance; Over-Sensing; Pacing Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM WAS OBSERVED TO EXHIBIT OVERSENSING NOISE AND PACING INHIBITION ON THE RIGHT VENTRICULAR (RV) LEAD. THE HEALTH CARE PROFESSIONAL (HCP) CALLED TECHNICAL SERVICES (TS) FOR FURTHER REVIEW OF THE PRESENTING ELECTROGRAM (EGM). IT WAS NOTED THAT THE PATIENT IS PACING DEPENDENT AND THAT THE RV LEAD IS A NON-BOSTON SCIENTIFIC LEAD. UPON REVIEW, TS NOTED THAT THE OVERSENSED RV LEAD NOISES APPEARED TO BE EITHER ELECTROMAGNETIC INTERFERENCE (EMI) OR MYOPOTENTIAL IN NATURE. TS ALSO NOTED A SUDDEN INTERMITTENT INCREASE IN RV LEAD PACING IMPEDANCE MEASUREMENTS FROM 480 TO 580 OHMS. TS DISCUSSED PROGRAMMING OPTIONS WITH THE HCP. AT THIS TIME, THE PACEMAKER AND NON-BOSTON SCIENTIFIC RV LEAD REMAIN IN SERVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20988968&pc=LWP | 2124215-2024-81024 | 2024-11-15 05:00 | 2024-12-20 05:00 | Malfunction | Signal Artifact/Noise; Low impedance | ACCOLADE MRI SR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM EXHIBITED LOW OUT OF RANGE IMPEDANCE MEASUREMENTS AND NOISY SIGNALS FOR ITS RIGHT VENTRICULAR (RV) LEAD. THE RV LEAD HAS BEEN IN SERVICE SINCE 1988. TS WAS CONSULTED AND RECOMMENDED FURTHER IN CLINIC EXAMINATION USING UNIPOLAR SETTING. THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM EXHIBITED LOW OUT OF RANGE IMPEDANCE MEASUREMENTS AND NOISY SIGNALS FOR ITS RIGHT VENTRICULAR (RV) LEAD. THE RV LEAD HAS BEEN IN SERVICE SINCE 1988. TS WAS CONSULTED AND RECOMMENDED FURTHER IN CLINIC EXAMINATION USING UNIPOLAR SETTING. THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21264129&pc=LWP | 2124215-2025-04394 | 2025-01-08 05:00 | 2025-01-29 05:00 | Malfunction | Signal Artifact/Noise; Over-Sensing | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM EXHIBITED NOISE OVERSENSING ON THE RIGHT VENTRICULAR (RV) CHANNEL. THE RV LEAD IS A NON-BOSTON SCIENTIFIC DEVICE. TECHNICAL SERVICES (TS) RECOMMENDED FURTHER EVALUATION OF THE LEAD, AS ABRASION OR INSULATION ISSUES WERE SUSPECTED. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21265382&pc=LWP | 2124215-2025-04450 | 2025-01-13 05:00 | 2025-01-29 05:00 | Malfunction | Signal Artifact/Noise; Over-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM EXHIBITED OVERSENSING OF NOISY SIGNALS FOR ITS RIGHT ATRIAL (RA) LEAD CHANNEL, SO AN ATRIAL TACHY RESPONSE (ATR) EPISODE WAS RECORDED. TECHNICAL SERVICES (TS) WAS CONSULTED AND DISCUSSED IT WAS ISOLATED AND LOOKED LIKE ELECTROMAGNETIC INTERFERENCE (EMI). THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21213947&pc=LWP | 2124215-2025-03056 | 2024-08-15 04:00 | 2025-01-22 05:00 | Malfunction | Signal Artifact/Noise; Over-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXPERIENCED A MINUTE VENTILATION (MV) SENSOR DISABLE DUE TO SIGNAL ARTIFACT MONITOR (SAM). TECHNICAL SERVICES (TS) WAS CONSULTED AND EXPLAINED THAT EXTERNAL NOISE WAS NOTED. THE DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21152274&pc=LWP | 2124215-2025-01318 | 2024-12-04 05:00 | 2025-01-14 05:00 | Malfunction | Signal Artifact/Noise; Over-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED OVERSENSING OF NOISE RESULTING IN INAPPROPRIATE ATRIAL TACHY RESPONSE (ATR) EPISODES. THE NOISE WAS FELT TO BE RELATED TO ELECTROMAGNETIC INTERFERENCE (EMI) FROM A POTENTIAL PROCEDURE. FURTHER REVIEW WAS RECOMMENDED. NO FURTHER ASSISTANCE WAS REQUESTED AT THIS TIME. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE. IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21123622&pc=LWP | 2124215-2025-00599 | 2021-11-21 05:00 | 2025-01-10 05:00 | Malfunction | Signal Artifact/Noise; Over-Sensing | ACCOLADE MRI EL DR | Discomfort; Presyncope | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER INAPPROPRIATELY STORED A SIGNAL ARTIFACT MONITOR (SAM) EPISODE AND DISABLED THE MINUTE VENTILATION (MV) SENSOR. TECHNICAL SERVICES (TS) REVIEWED THE DATA, AND IT WAS SUSPECTED THAT THE SAM EPISODE WAS STORED DUE TO ELECTROMAGNETIC INTERFERENCE (EMI) OVERSENSING DURING A PROCEDURE. THE HEALTH CARE PROFESSIONAL (HCP) PROGRAMMED MV BACK ON, AND THE PATIENT HAD FELT SOME DISCOMFORT AND SYMPTOMS OF PRESYNCOPE, HOWEVER, IT WAS NOT CONCLUSIVE WHETHER IT WAS DEVICE RELATED. TS DISCUSSED PROGRAMMING OPTIONS. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21105169&pc=LWP | 2124215-2024-83162 | 2024-08-14 04:00 | 2025-01-08 05:00 | Malfunction | Signal Artifact/Noise; Over-Sensing | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED NOISE ON THE RIGHT ATRIAL (RA) LEAD. THE NOISE WAS OVERSENSED, RESULTING IN AN INAPPROPRIATE ATRIAL TACHY RESPONSE (ATR) EPISODE TO BE STORED. ADDITIONALLY, THE NOISE CAUSED A SIGNAL ARTIFACT MONITOR (SAM) EPISODE AND CAUSED THE MINUTE VENTILATION (MV) SENSOR TO BE DISABLED. ALERTS WERE ISSUED IN THE REMOTE MONITORING SYSTEM AND AN EMAIL WAS SENT TO THE CLINIC FOR NOTIFICATION. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THE DEVICE AND ASSOCIATED RA LEAD REMAIN IMPLANTED AND IN SERVICE. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21048184&pc=LWP | 2124215-2024-82614 | 2024-10-15 04:00 | 2024-12-31 05:00 | Malfunction | Signal Artifact/Noise; Over-Sensing | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM EXHIBITED NOISE. BOSTON SCIENTIFIC TECHNICAL SERVICES (TS) DISCUSSED POSSIBLE LEAD TO LEAD ABRASION. THE PATIENT REPORTED TO HAVE BEEN LIFTING WEIGHTS. THE NOISE WAS OVERSENSED. THE NOISE HAS BEEN HISTORICALLY REPRODUCIBLE WITH ISOMETRIC MOVEMENT. THIS PACEMAKER SYSTEM REMAINS IN SERVICE AND WILL CONTINUE TO BE MONITORED. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21035604&pc=LWP | 2124215-2024-82356 | 2024-12-05 05:00 | 2024-12-30 05:00 | Malfunction | Signal Artifact/Noise; Over-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THERE WERE CONCERNS ABOUT MINUTE VENTILATION (MV) CHOP ON STORED VENTRICULAR TACHYCARDIA (VT) EPISODES ON THIS PACEMAKER. TECHNICAL SERVICES (TS) REVIEWED THE EVENTS AND NOTED NON-PHYSIOLOGICAL NOISY SIGNALS ON BOTH THE RIGHT ATRIAL (RA) AND RIGHT VENTRICULAR (RV) CHANNELS. THE RV DEFLECTIONS WERE SENSED BY THE DEVICE. HOWEVER, MV CHOP WAS DISCARDED AS A CAUSE SINCE IT WOULD HAVE MORE DEFLECTIONS AND BE FASTER. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20921191&pc=LWP | 2124215-2024-78828 | 2024-11-27 05:00 | 2024-12-12 05:00 | Malfunction | Signal Artifact/Noise; Over-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THE SIGNAL ARTIFACT MONITOR (SAM) OF THE PACEMAKER DISABLED THE MINUTE VENTILATION (MV) FEATURE DUE TO SENSING OF NOISE. THE NOISE WAS PRESENT ON BOTH THE ATRIAL AND VENTRICULAR CHANNELS AND THERE WAS A FALSE EPISODE OF NON-SUSTAINED VENTRICULAR TACHYCARDIA DUE TO THE OVERSENSING. FURTHER REVIEW OF THE DEVICE SETTINGS WAS RECOMMENDED. THE PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21107949&pc=LWP | 2124215-2025-00359 | 2024-09-12 04:00 | 2025-01-08 05:00 | Malfunction | Signal Artifact/Noise; Over-Sensing; Device Sensing Problem | ACCOLADE EL DR | No Clinical Signs, Symptoms or Conditions | N970003 | 1 | Event Description: IT WAS REPORTED THAT ELECTROGRAM (EGM) REVIEW OF THIS PACEMAKER SYSTEM REVEALED FARFIELD R-WAVE OVERSENSING AND RIGHT ATRIAL (RA) LEAD NOISE WITH OVERSENSING. ADDITIONALLY, THERE WERE ONLY FOUR SUCCESSFUL RIGHT ATRIAL AUTOMATIC THRESHOLD (RAAT) TESTS OVER THE PAST YEAR. IT WAS NOTED THAT IMPEDANCE AND SENSING WERE VERY STABLE. TECHNICAL SERVICES (TS) DISCUSSED TROUBLESHOOTING OPTIONS AND PROGRAMMING CONSIDERATIONS. THIS PACEMAKER REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21101360&pc=LWP | 2124215-2025-00186 | 2023-12-01 05:00 | 2025-01-08 05:00 | Malfunction | Signal Artifact/Noise; Over-Sensing; Failure to Read Input Signal | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER DEVICE EXHIBITED OVERSENSING AND NOISE ON NON-BOSTON SCIENTIFIC RIGHT VENTRICULAR (RV) CHANNEL. UPON REVIEW, IT WAS NOTED THAT THERE WAS NO NOISE OBSERVED IN NOVEMBER 2023, DURING ATRIAL AUTOMATIC THRESHOLD TEST. HOWEVER, THE NOISE OBSERVED CURRENTLY ON BOTH RA AND RV ELECTROGRAM (EGM) AND THIS WAS AFTER THE LAST CASE IN DECEMBER 2023. THE RIGHT VENTRICULAR (RV) LEAD ALSO SHOWED FLUCTUATION IN SENSING, IMPEDANCE AND THRESHOLD VALUES. THE INTRINSIC AMPLITUDE WAS ALSO OUT OF RANGE FOR RIGHT ATRIAL (RA) LEAD. TECHNICAL SERVICES (TS) RECOMMENDED TO HAVE THE PATIENT SEEN IN CLINIC FOR TROUBLESHOOTING AND PERFORM FULL CHEST X-RAY IMAGING. THIS DEVICE AND NON-BOSTON SCIENTIFIC LEAD CURRENTLY REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20984131&pc=LWP | 2124215-2024-80849 | 2024-12-03 05:00 | 2024-12-19 05:00 | Malfunction | Signal Artifact/Noise; Over-Sensing; Failure to Read Input Signal | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED NOISE ON THE RIGHT ATRIAL (RA) CHANNEL. TS REVIEWED THE RECORDED ATRIAL TACHY RESPONSE (ATR) EPISODE AND INDICATED THAT THE NOISE MAY HAVE BEEN DUE TO ELECTROMAGNETIC INTERFERENCE (EMI) DURING A PROCEDURE. IT WAS NOTED THAT THE DEVICE APPEARED TO BE SENSING AND CAPTURING APPROPRIATELY. ADDITIONALLY, TS NOTED THAT THERE WERE CONCERNING LOW P WAVE AMPLITUDES AND RECOMMENDED THAT THE PATIENT FOLLOW-UP IN CLINIC WITH A LOCAL BOSTON SCIENTIFIC REPRESENTATIVE. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS RA LEAD REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20965788&pc=LWP | 2124215-2024-80249 | 2024-12-10 05:00 | 2024-12-18 05:00 | Malfunction | Signal Artifact/Noise; Over-Sensing; Failure to Read Input Signal | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER RECORDED OUT OF RANGE ATRIAL INTRINSIC AMPLITUDE MEASUREMENTS RESULTING IN AN ALERT IN THE REMOTE HOME MONITORING SYSTEM. REVIEW OF STORED DEVICE DATA IN THE REMOTE HOME MONITORING SYSTEM NOTED THIS PACEMAKER AND RIGHT ATRIAL (RA) LEAD EXHIBITED OVERSENSING OF NOISE, WHICH, RESULTED IN AN INAPPROPRIATE ATRIAL TACHY RESPONSE (ATR) MODE SWITCH. FURTHER REVIEW WAS RECOMMENDED. NO FURTHER ASSISTANCE WAS REQUESTED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE. IF PERTINENT INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21144818&pc=LWP | 2124215-2025-01089 | 2024-09-26 04:00 | 2025-01-13 05:00 | Malfunction | Signal Artifact/Noise; Over-Sensing; Off-Label Use; Failure to Read Input Signal; Use of Device Problem | ACCOLADE MRI SR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED OVERSENSING OF MYOPOTENTIAL SIGNALS. ADDITIONALLY, IT WAS DISCUSSED THAT THIS PATIENT HAS HAD A HISTORY OF LOW R WAVES SINCE THIS SYSTEM WAS IMPLANTED. THE RIGHT VENTRICULAR AUTOMATIC THRESHOLD (RVAT) TEST WAS UNSUCCESSFUL DUE TO LOW EVOKE RESPONSE. IT WAS NOTED THAT THIS SYSTEM IS IMPLANTED WITH A NON BOSTON SCIENTIFIC LEAD IN THE LEFT BUNDLE BRANCH. FURTHER DISCUSSION WITH THE PHYSICIAN AND TROUBLESHOOTING OPTIONS WERE RECOMMENDED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. AT THIS TIME, THIS DEVICE REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21201314&pc=LWP | 2124215-2025-02712 | 2025-01-04 05:00 | 2025-01-21 05:00 | Malfunction | Signal Artifact/Noise; Over-Sensing; Pacing Problem | ACCOLADE MRI EL DR | Asystole | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM TRIGGERED A LEAD SAFETY SWITCH (LSS) DUE TO HIGH OUT-OF-RANGE PACING IMPEDANCES IN THE RIGHT VENTRICULAR (RV) LEAD, MEASURING ABOVE 2000 OHMS. THE INCREASE IN PACING IMPEDANCES WAS NOTED TO BE GRADUAL. ADDITIONALLY, THE SYSTEM EXHIBITED NOISE OVERSENSING AND PACING INHIBITION, RESULTING IN AN ASYSTOLE LASTING MORE THAN TWO SECONDS. TECHNICAL SERVICES (TS) REVIEWED THE EVENT AND RECOMMENDED FURTHER EVALUATION OF THE LEAD, INCLUDING ISOMETRIC EXERCISES AND POCKET MANIPULATIONS, AS LEAD CALCIFICATION WAS SUSPECTED. TS ALSO PROVIDED REPROGRAMMING OPTIONS. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21184477&pc=LWP | 2124215-2025-02154 | 2025-01-09 05:00 | 2025-01-17 05:00 | Malfunction | Signal Artifact/Noise; Over-Sensing; Pacing Problem | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions |  | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED OVERSENSING OF MYOPOTENTIAL NOISE RESULTING IN PACING INHIBITION. THE DEVICE WAS PROGRAMMED TO UNIPOLAR TO MITIGATE FURTHER OVERSENSING. THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21152273&pc=LWP | 2124215-2025-01324 | 2025-01-03 05:00 | 2025-01-14 05:00 | Malfunction | Signal Artifact/Noise; Over-Sensing; Pacing Problem | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT A SCHEDULED REMOTE MONITORING TRANSMISSION REVIEW SHOWED RECORDED ATRIAL TACHY RESPONSE (ATR) EVENTS WHICH SHOWED NOISE OVERSENSING ON THE ATRIAL CHANNEL RESULTING IN ATRIAL INHIBITION OF PACE. THE DEVICE REMAINS IMPLANTED TO DATE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21114411&pc=LWP | 2124215-2025-00497 | 2024-11-25 05:00 | 2025-01-09 05:00 | Malfunction | Signal Artifact/Noise; Over-Sensing; Pacing Problem | ACCOLADE MRI EL DR | Asystole | P150012 | 1 | Event Description: IT WAS REPORTED THAT BASELINE NOISE OVERSENSING WAS PRESENT ON THE RIGHT VENTRICULAR (RV) LEAD CHANNEL DURING A STORED ATRIAL TACHY RESPONSE (ATR) ON THIS PACEMAKER. THE NOISE OVERSENSING RESULTED IN PACING INHIBITION UP TO 2.5 SECONDS. THE PACEMAKER IS PROGRAMMED WITH AUTOMATIC GAIN CONTROL AT 0.6 MV. TECHNICAL SERVICES (TS) RECOMMENDED IN OFFICE FOLLOW UP WITH TROUBLESHOOTING, AS WELL AS REPROGRAMMING TO FIXED SENSING. THIS PACEMAKER AND RIGHT VENTRICULAR (RV) LEAD REMAIN IN-SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21105773&pc=LWP | 2124215-2024-83202 | 2023-05-06 04:00 | 2025-01-08 05:00 | Malfunction | Signal Artifact/Noise; Over-Sensing; Pacing Problem | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT REVIEW OF STORED ELECTROGRAMS (EGMS) SHOWED INTERMITTENT NOISE AND OVERSENSING ON THE ATRIAL AND RIGHT VENTRICULAR (RV) CHANNELS. FOUR SECONDS OF PACING INHIBITION OCCURRED DUE TO THE OVERSENSED NOISE. IT WAS NOTED THAT THE EVENT WAS WHILE THIS PATIENT WAS UNDERGOING HIP SURGERY. THE SYSTEM REMAINS IN SERVICE AND NO ADVERSE PATIENT EFFECTS WERE REPORTED. Manufacturer Narrative: AS NO FURTHER INFORMATION CONCERNING THIS REPORT IS EXPECTED, OUR INVESTIGATION IS COMPLETE. THIS INVESTIGATION WILL BE UPDATED SHOULD FURTHER INFORMATION BE PROVIDED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20948689&pc=LWP | 2124215-2024-79557 | 2024-12-05 05:00 | 2024-12-17 05:00 | Malfunction | Signal Artifact/Noise; Pacing Problem | ACCOLADE MRI DR | Asystole | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXHIBITING NOISE SIGNALS AND PACING INHIBITION ON THE RIGHT VENTRICULAR (RV) LEAD. THE PATIENT EXPERIENCED AN ASYSTOLE. THE LEAD WILL BE REVISED IN THE FUTURE. THE DEVICE REMAINS IN SERVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED.Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXHIBITING NOISE SIGNALS AND PACING INHIBITION ON THE RIGHT VENTRICULAR (RV) LEAD. THE PATIENT EXPERIENCED AN ASYSTOLE. THE LEAD WILL BE REVISED IN THE FUTURE. THE DEVICE REMAINS IN SERVICE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION WAS OBTAINED, THE DEVICE EXPERIENCED OVERSENSING.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21242075&pc=LWP | 2124215-2025-03744 | 2025-01-13 05:00 | 2025-01-27 05:00 | Malfunction | Signal Artifact/Noise; Under-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM DISPLAYED YELLOW ALERT FOR RIGHT VENTRICULAR AUTOMATIC THRESHOLD (RVAT) DETECTED AS GREATER THAN PROGRAMMED AMPLITUDE OR SUSPENDED. TECHNICAL SERVICES (TS) WAS CONSULTED AND NOTED UNDERSENSING OF ATRIAL FIBRILLATION (AF) AND INTRINSIC AMPLITUDE OUT OF RANGE. THE DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20972687&pc=LWP | 2124215-2024-80502 | 2024-11-24 05:00 | 2024-12-18 05:00 | Malfunction | Telemetry Discrepancy | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER SYSTEM WAS UNABLE TO BE INTERROGATED WITH A PROGRAMMER. THE HEALTH CARE PROFESSIONAL (HCP) CALLED TECHNICAL SERVICES (TS) NOTING PROGRAMMER WOULD NOT READ THE COMPATIBLE DEVICE AND REQUESTED TROUBLESHOOTING EFFORTS. TS REVIEWED THE PROGRAMMER SETTINGS AND DETERMINED CAUSE TO BE DUE TO A PROGRAMMER SOFTWARE ISSUE. TROUBLESHOOTING EFFORTS WERE ATTEMPTED BUT WERE UNSUCCESSFUL. TS DISCUSSED POSSIBLE NEXT STEPS WITH THE HCP AND RECOMMENDED THE LOCAL FIELD REPRESENTATIVE TO BE CONTACTED FOR FURTHER ASSISTANCE. AT THIS TIME, THE PROGRAMMER AND PACEMAKER REMAIN IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21267227&pc=LWP | 2124215-2025-04503 | 2025-01-05 05:00 | 2025-01-29 05:00 | Malfunction | Under-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTED PACEMAKER ALERTED FOR AN ATRIAL ARRHYTHMIA BURDEN. THERE WAS OCCASIONAL ATRIAL UNDER SENSING OBSERVED AT PROGRAMMED SENSITIVITY FIXED 1.0MV (MILLIVOLTS). IT WAS RECOMMENDED TO REVIEW PROGRAMMED PARAMETERS, AND THE ATRIAL ARRHYTHMIA BURDEN ALERT WAS INCREASED TO GREATER THAN 24 HOURS FOR FUTURE EPISODES. THE BATTERY STATUS IS NORMAL AND THE OTHER LEAD MEASUREMENTS ARE SATISFACTORY. THE PACEMAKER REMAINS IN SERVICE AND THERE WAS NO ADVERSE PATIENT EFFECTS REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21267773&pc=LWP | 2124215-2025-04536 | 2025-01-12 05:00 | 2025-01-29 05:00 | Malfunction | Under-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER ALERTED FOR AN ATRIAL ARRHYTHMIA BURDEN GREATER THAN 24 HOURS. THE PRESENTING ELECTROGRAM (EGM) SHOWS AN ONGOING ATRIAL ARRHYTHMIA WITH OCCASIONAL UNDERSENSING OF ATRIAL SIGNALS OBSERVED. THE ATRIAL INTRINSIC AMPLITUDE RANGES FROM 0.7MV (MILLIVOLTS) TO 3.8MV ON TREND WITH THE SENSITIVITY PROGRAMMED TO AUTOMATIC GAIN CONTROL (AGC) AT 0.25MV. THE BATTERY LONGEVITY IS NORMAL, AND THE LEAD MEASUREMENTS ARE STABLE. FURTHER REVIEW WAS RECOMMENDED. THE DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21249268&pc=LWP | 2124215-2025-03906 | 2025-01-13 05:00 | 2025-01-27 05:00 | Malfunction | Under-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXHIBITING ATRIAL UNDERSENSING. THE DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21233149&pc=LWP | 2124215-2025-03530 | 2024-08-26 04:00 | 2025-01-24 05:00 | Malfunction | Under-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THE ATRIAL ARRHYTHMIA BURDEN ALERT WAS TRIGGERED AND ELECTROGRAM (EGM)S SHOWED UNDERSENSING OF AN ATRIAL ARRHYTHMIA. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THE DEVICE REMAINS IMPLANTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21203047&pc=LWP | 2124215-2025-02733 | 2025-01-07 05:00 | 2025-01-21 05:00 | Malfunction | Under-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXHIBITING UNDERSENSING IN WHICH ATRIAL FLUTTER BEATS WERE FALLING INTO BLANKING FROM VENTRICULAR PACING AND AS A RESULT WOULD NOT MODE SWITCH. TECHNICAL SERVICES (TS) RECOMMENDED PROGRAMMING CHANGES TO SHORTEN THE ATRIAL BLANKING PERIODS. FOLLOWING THESE CHANGES THE DEVICE CORRECTLY MODE SWITCHED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS PACEMAKER REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21204258&pc=LWP | 2124215-2025-02808 | 2024-12-30 05:00 | 2025-01-21 05:00 | Malfunction | Under-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER ALERTED FOR AN ATRIAL ARRHYTHMIA BURDEN. THE STORED ELECTROGRAMS (EGMS) SHOWED ATRIAL TACHY RESPONSE (ATR) WITH UNDERSENSING OF ATRIAL FIBRILLATION (AF) OBSERVED. BATTERY LONGEVITY IS NORMAL AND OTHER LEAD MEASUREMENTS ARE STABLE. AT THIS TIME, THE DEVICE REMAINS IN-SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21192541&pc=LWP | 2124215-2025-02429 | 2024-12-23 05:00 | 2025-01-20 05:00 | Malfunction | Under-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER DEVICE EXHIBITED INTERMITTENT UNDERSENSING P WAVE 0.5 MILLI VOLTS AND PROGRAMMED 0.75 MILLI VOLTS. HEALTH CARE PROFESSIONAL (HCP) INQUIRED FOR A REPROGRAMMING CHANGE. AS OF THIS TIME, THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21151742&pc=LWP | 2124215-2025-01303 | 2024-12-23 05:00 | 2025-01-14 05:00 | Malfunction | Under-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT PRESENTING ELECTROGRAM (EGM) SHOWED ATRIAL UNDERSENSING. NO CHANGES WERE MADE. THE DEVICE REMAINS IMPLANTED. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21123759&pc=LWP | 2124215-2025-00700 | 2022-05-11 04:00 | 2025-01-10 05:00 | Malfunction | Under-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER EXHIBITED UNDERSENSING. BOSTON SCIENTIFIC TECHNICAL SERVICES (TS) OBSERVED UNDERSENSING AS WELL. THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21129683&pc=LWP | 2124215-2025-00829 | 2024-12-12 05:00 | 2025-01-10 05:00 | Malfunction | Under-Sensing | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER DEVICE HAD EXHIBITED ATRIAL UNDERSENSING DESPITE PROGRAMMING AT THE MOST SENSITIVE. BOSTON SCIENTIFIC TECHNICAL SERVICES (TS) DISCUSSED THAT THE MOST RECENT LEAD TESTS ARE WITHIN NORMAL LIMITS AND THE PRESENTING ELECTROGRAM SHOWS THE DEVICE OPERATING AS PROGRAMMED. THE IMPLANTABLE PACEMAKER DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21054919&pc=LWP | 2124215-2025-00023 | 2024-12-02 05:00 | 2025-01-02 05:00 | Malfunction | Under-Sensing | ACCOLADE MRI DR | Unspecified Heart Problem | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS IMPLANTABLE PACEMAKER DEVICE EXHIBITED AN ATRIAL UNDERSENSING WHICH WAS NOTED ON THE RYTHMIQ EVENTS. ALSO, PATIENT EXPERIENCED PALPITATIONS. BOSTON SCIENTIFIC TECHNICAL SERVICES (TS) REVIEWED THE EVENTS AND INDICATED THAT THIS DEVICE APPEARED TO BE FUNCTIONING AS PROGRAMMED. AS OF THIS TIME, THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21028293&pc=LWP | 2124215-2024-82199 | 2024-11-12 05:00 | 2024-12-27 05:00 | Malfunction | Under-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER EXHIBITED UNDERSENSING OF ATRIAL FIBRILLATION (AF). TECHNICAL SERVICES (TS) WAS CONSULTED AND NOTED THAT IN THE STORED ATRIAL TACHY RESPONSE (ATR) EPISODES, AND IN THE LOGBOOK REPORT, THE DEVICE WAS ABLE TO MODE SWITCH BUT DOES APPEAR TO DROP OUT OF MODE SWITCH FROM TIME TO TIME. A FOLLOW UP WITH CARDIOLOGY WAS RECOMMENDED. NO ADVERSE PATIENT EFFECTS WERE REPORTED. THIS DEVICE REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21018993&pc=LWP | 2124215-2024-81941 | 2024-11-26 05:00 | 2024-12-26 05:00 | Malfunction | Under-Sensing | ACCOLADE MRI DR | Arrhythmia; Dizziness | P150012 | 1 | Event Description: IT WAS REPORTED THAT THE PATIENT WITH THIS IMPLANTABLE PACEMAKER DEVICE EXPERIENCED LIGHT HEADEDNESS. BOSTON SCIENTIFIC TECHNICAL SERVICES (TS) ASSESSED THE DATA TRANSMISSION AND INDICATED THAT NON-SUSTAINED VENTRICULAR TACHYCARDIA (NSVT) WAS REAL OR TRUE WHICH COULD BE CONTRIBUTING TO THE SYMPTOMS. ARRHYTHMIA WAS ALSO NOTED WHICH SHOWED AN INTERMITTENT UNDERSENSING PREMATURE VENTRICULAR CONTRACTION (PVC) LIKELY DUE TO AUTOMATIC GAIN CONTROL (AGC) SENSING METHOD AT 1.5 MILLI VOLTS. TS SUGGESTED TO CONSIDER MEASURING THE UNDERSENSED PVCS AND PROGRAM AN APPROPRIATE FIXED SENSING VALUE. AS OF THIS TIME, THIS DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=21019927&pc=LWP | 2124215-2024-81958 | 2024-12-16 05:00 | 2024-12-26 05:00 | Malfunction | Under-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS SUSPECTED OF EXHIBITING UNDERSENSING. TECHNICAL SERVICES (TS) RECOMMENDED REPROGRAMMING OPTIONS. NO ADVERSE PATIENT EFFECTS WERE REPORTED. AT THIS TIME, THIS DEVICE REMAINS IN SERVICE.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20985140&pc=LWP | 2124215-2024-80902 | 2024-08-01 04:00 | 2024-12-20 05:00 | Malfunction | Under-Sensing | ACCOLADE MRI DR | No Clinical Signs, Symptoms or Conditions | P150012/S000 | 1 | Event Description: IT WAS REPORTED THAT A SCHEDULED TRANSMISSION REVIEW ON THIS IMPLANTABLE PACEMAKER SHOWED ATRIAL TACHY RESPONSE (ATR) WITH ATRIAL UNDERSENSING OF ARRYTHMIA. THE BATTERY LONGEVITY IS NORMAL, AND THE LEAD MEASUREMENTS ARE STABLE. THE DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
| https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfMAUDE/Detail.CFM?MDRFOI__ID=20956406&pc=LWP | 2124215-2024-80005 | 2024-12-05 05:00 | 2024-12-17 05:00 | Malfunction | Under-Sensing | ACCOLADE MRI EL DR | No Clinical Signs, Symptoms or Conditions | P150012 | 1 | Event Description: IT WAS REPORTED THAT THIS PACEMAKER WAS EXHIBITING FUNCTIONAL UNDERSENSING. THE DEVICE PRESENTED LOSS OF CAPTURE DUE TO THE VENTRICLE BEING IN REFRACTORY. THE RIGHT VENTRICULAR (RV) LEAD PACING WAS PRESENT DUE TO INTRINSIC BEATS. THE DEVICE REMAINS IN SERVICE. NO ADVERSE PATIENT EFFECTS WERE REPORTED.
MDR data alone cannot be used to establish rates of events | evaluate a change in event rates over time or compare event rates between devices. The number of reports cannot be interpreted or used in isolation to reach conclusions about the existence | severity | or frequency of problems associated with devices. Confirming whether a device actually caused a specific event can be difficult based solely on information provided in a given report. Establishing a cause-and-effect relationship is especially difficult if circumstances surrounding the event have not been verified or if the device in question has not been directly evaluated.
If a report contains trade secret or confidential business information | that text is replaced by "(b)(4)". If a report contains personnel or medical files information | that text is replaced by "(b)(6)". The designations "(b)(4)" and "(b)(6)" refer to the exemptions in the FOIA. For example | "(b)(4)" may be found in place of the product's composition and "(b)(6)" may be found in place of a patient's age. MAUDE is updated monthly and the search page reflects the date of the most recent update. The FDA seeks to include all reports received prior to the update but the inclusion of some reports may be delayed.
Manufacturers and importers must submit reports when they become aware of information that reasonably suggests that one of their marketed devices may have caused or contributed to a death or serious injury or has malfunctioned and the malfunction of the device or a similar device that they market would be likely to cause or contribute to a death or serious injury if the malfunction were to recur. Manufacturers must send reports of such deaths | serious injuries and malfunctions to the FDA. Importers must send reports of deaths and serious injuries to the FDA and the manufacturer | and reports of malfunctions to the manufacturer.